INTRODUCTION

Dentistry has changed tremendously over the past decade to the benefit of both the clinician and the patient. New materials and technologies have improved the efficiency and predictability of treatment for clinicians. The computer has increased its influence in dental field with digital radiography and photography, and CAD-CAM. New techniques and philosophies have also benefited our patients by providing alternatives to aggressively prepared teeth and the use of aesthetic-compromising metal.

One technology that has become increasingly utilized in clinical dentistry is that of the laser. Lasers were discovered in 1960 by Theodore H. Maiman and are used today in a great variety of applications, from compact disc players to laser eye surgery. Laser devices produce a very strong, coherent, unidirectional, and monochromatic light beam, which can be controlled to perform special functions1.

Initially introduced as an alternative to the traditional halogen curing light, the laser now has become the instrument of choice, in many applications, for both periodontal and restorative care. Although the laser offers many advantages over other modalities of treatment, probably the greatest impact it has made, is its ability to be used for both hard and soft tissue, often  without the need for anesthesia. The US Food and Drug

Administration has approved the use of some types of lasers for medical and dental procedures, of which the carbon dioxide, neodymium:yttrium-aluminum-garnet (Nd:YAG), and argon lasers are among the most popular approved laser types.

The word ‘Laser’ is an acronym for Light Amplification by Stimulated Emission of Radiation. The energy generated by the laser is in or near the optical portion of the electromagnetic spectrum which is non-ionizing. It does not produce the same harmful effects attributed to x-rays and other forms of ionizing radiation. Energy is amplified to extremely high intensity by an atomic process called stimulated emission.

The term “radiation” is often misinterpreted since it is used to describe radioactive materials or ionizing radiation. The use of the word in this context, however, refers to an energy transfer. Energy moves from one location to another by conduction, convection, and radiation. The colour of laser light is normally expressed in terms of the laser’s wavelength. The most common unit used in expressing laser’s wavelength is a nanometer (nm). There are one billion nanometers in one meter (1 nm = 1 X 10-9 m). Laser light is non-ionizing and includes ultra-violet (100-400nm), visible (400-700nm), and infrared (700nm-1mm)2.

The future of dental lasers is bright, with some of the newest ongoing research addressing the use of dental lasers for guided tissue regeneration, connective tissue attachment and tissue welding. Lasers of the future may make incisions, reflect flaps,

perform osseous surgery and weld flaps into position incorporating multiple wavelengths into a single unit.

Dentistry has entered an exciting era of high technology with lasers offering the dentist not only a window, but a door into this high-tech, rewarding and potentially profitable arena3.

WHEN

HISTORY OF LASERS

The principle of the laser was first known in 1917, when physicist Albert Einstein described the theory of stimulated emission. However, it was not until the late 1940s that engineers began to utilize this principle for practical purposes. At the onset of 1950’s several different engineers were working towards  harnessing of energy using the principal of stimulated emission.

At the University of Columbia was Charles Townes, at the University of Maryland was Joseph Weber and at the Lebedev Laboratories in Moscow were Alexander Prokhorov and Nikolai G Basov who worked towards the creation of what was termed a “MASER” (Microwave Amplification by the Stimulated Emission of Radiation), a device that amplified microwaves by passing it through ammonia gas. It was realized that only a fraction of the incident energy was converted into maser energy, the greater emission being in the form of heat ; the output power of the early masers was of the order of a few micro-watts and hence found use in microwave communication systems.

Townes and the other engineers believed it to be possible to create an optical maser, a device for creating powerful beams of light using higher frequency energy to stimulate what was to become termed the lasing medium. Despite the pioneering work of Townes and Prokhorov, it was Theodore Maiman who on May 16th 1960 invented the

first Laser at Hughes Research Laboratories in California using a lasing medium of ruby that was stimulated using high energy flashes of intense light.

When the first working laser was reported in 1960, it was considered a technology without a purpose and was described as "a solution looking for a problem." Before long, the laser's distinctive qualities, its ability to generate an intense, very narrow beam of light of a single wavelength were being harnessed for science, technology and medicine. Today, lasers are everywhere, from research laboratories at the cutting edge of quantum physics to medical clinics, supermarket checkouts and telephone networks4.

Maiman had exposed an extracted tooth to his ruby laser in 1960. The first application of a laser to dental tissue was reported by Goldman et al. & Stern and Sognnaes. Each of these articles described the effects of the ruby laser on enamel and dentin. The possibilities for laser use in dentistry did not occur until it was introduced by the American Dental Laser for commertial use in the year 1989. This laser, using an active medium of Nd:YAG, emitted pulsed light, which was developed and marketed by an American dentist, Dr Terry Myers. Though low-powered, and its emission wavelength inappropriate for use on dental hard tissue, it still gained popularity among dentists. This laser was first sold in the United Kingdom in 1990.Other laser wavelengths, using machines that were already in use in medicine and surgery and only slightly modified, became available for dental use in the early 1990s.

Being predominately argon, Nd:YAG, CO2 and semiconductor diodes, all these lasers failed to address a growing need among dentists and patients for a laser that would ablate dental hard tissue. In 1989, the experimental work by Keller and Hibst using a pulsed erbium YAG (2,940nm) laser, demonstrated its effectiveness in cutting enamel, dentine and bone. This laser became commercially available in the United Kingdom in 1995, shortly followed by a similar Er, Cr:YSGG (erbium chromium : yttrium scandium gallium garnet) laser in 1997, which amounted to a laser armamentarium that would address the surgical needs of clinical dentistry in general practice5.

COMPONENTS OF A TYPICAL LASER

1. ACTIVE MEDIUM

A material, either naturally occurring or man-made that when stimulated, emits laser light. This material may be a solid, liquid or gas. The first ‘dental’ laser used a crystal of neodymium-doped yttrium aluminium garnet (Nd:Yag) as its active medium. ‘YAG’ is a complex crystal with the chemical composition Y3Al5O12. During crystal growth, 1% neodymium (Nd3+) ions are doped into the YAG crystal.

Other lasers of significance in dentistry use rare earth and other metal ions within a ‘doped’ YAG crystal lattice, eg - Erbium (Er:YAG) and holmium (Ho:YAG), together with another erbium and chromium-doped garnet of yttrium, scandium and gallium (Er, Cr:YSGG).

The active medium is positioned within the laser cavity, an internally-polished tube, with mirrors co-axially positioned at each end and surrounded by the external energizing input, or pumping mechanism. The ‘active medium’, eg CO2, Nd:YAG, defines the type of laser and the emission wavelength of the laser (10,600nm and 1,064nm respectively). Atoms of the active medium are absorbed by the process of light emission.

2. PUMPING MECHANISM

This represents a man-made source of primary energy that excites the active medium. This is usually a light source, either a flash light or arc-light, but can be a diode laser unit or an electromagnetic coil. Energy from this primary source is absorbed by the active medium, resulting in the production of laser light.

The dynamics of incident energy with time has a fundamental bearing on the emission mode characteristics of a given laser. A continuous-feed electrical discharge will result in a similar continuous feed of laser light emission.

3. OPTICAL RESONATOR

Laser light produced by the stimulated active medium is bounced back and forth through the axis of the laser cavity, using two mirrors placed at either end, thus amplifying the power. The distal mirror is partly transmissive, so that at a given energy density, laser light will escape to be transmitted to the target tissue.

4. DELIVERY SYSTEM

Depending upon the emitted wavelength, the delivery system may be a quartz fibre-optic, a flexible hollow waveguide, an articulated arm (incorporating mirrors), or a hand-piece containing the laser unit (at present only for low-powered lasers).

Early attempts to produce delivery systems relied upon the use of fixed mirror and/or lens apparatus. It was soon apparent that the use of a fine, silica quartz fibre-optic cable maximised the feasibility for medical and dental lasers to reach their target site. However, the suitability of this delivery system is conditional upon the emission wavelength being poorly absorbed by water (hydroxyl groups), present in the quartz fibre.

Therefore shorter wavelengths (Er,Cr:YSGG, Er:YAG and carbon dioxide) give rise to severe power losses through quartz fibre and hence require alternative delivery systems.

Examples of such alternatives are articulated arms incorporating internal mirrors and prisms, and hollow waveguides, where the light is reflected along internally-polished tubes. Newer, water-free fibre compounds, eg : zirconium fluoride, are being developed to overcome this problem.

5. COOLING SYSTEM

Heat production is a by-product of laser light propagation. It increases with the power output of the laser, hence; with heavy-duty tissue cutting lasers, the cooling system represents the bulkiest component. Co-axial coolant systems may be air- or water-assisted.

6. CONTROL PANEL

This allows variation in power output with time, above that defined by the pumping mechanism frequency. Other facilities may allow wavelength change (multi-laser) instruments) and print-out of delivered laser energy during clinical use6.

DELIVERY SYSTEM

The field of lasers in general practice essentially began with the introduction of the American Dental Laser (Birmingham, Michigan) dLase 300 neodymium:yttrium-aluminium-garnet (Nd:YAG) laser in 1990. Before the introduction of this instrument, most dental lasers used bulky articulated arms for their delivery systems. These articulated arms were not conductive to the practice of general dentistry, owing to the long learning curve needed to master their use and the difficulty of delivering the laser energy to the entire oral cavity.

Articulated arm delivery systems consist of a series of rigid hollow tubes with mirrors at each joint (called a knuckle) that reflect the energy down the length of the tube. These joints exist to allow the delivery arm to be bent and configured in such a way as to bring the handpiece close to the target tissue. The laser energy exits the tube through a handpiece.

To transmit the laser energy efficiently, the mirrors at each knuckle must be aligned precisely. A misalignment of the mirrors could cause a drop-off in the amount of energy transmitted to the handpiece. The mirrors could go out of alignment through the normal use of moving the articulated arm for each new procedure or if the laser is moved from one treatment room to the other.

Articulated arm delivery systems are noncontact systems (ie, the handpiece or its attachments do not come into contact with the target tissue). Dentists are familiar with contact technology. The fissure bur contacts the enamel during tooth preparation. The curet contacts the root surface during scaling and root planning. Using a technology in which there is no contact between the instrument and target tissue can be challenging at first. This is one major reason for a longer learning curve with these instruments compared with contact technology instruments.

The American Dental Laser dLase Nd:YAG system was the first such instrument to use a fiberoptic delivery system. This fiberoptic technology allows for contact with the target tissue. The fiberoptic cables are attached to a small handpiece similar in size to a dental turbine and are available in sizes ranging from 200µm to 1000µm in diameter. They also are relatively flexible which allows easy transmission of the laser energy throughout the oral cavity, including periodontal pockets. Fiberoptic delivery and articulated arm systems are not the only two delivery systems currently in the market.

One manufacturer has developed a hollow waveguide delivery system. In contrast to an articulated arm system, this waveguide is a single long, semiflexible tube, without knuckles or mirrors. The laser energy is transmitted along the reflective inner lumen of this tube and exits through a handpiece at the end of the tube.

This handpiece comes with various attachments that the dentist may select, depending on the procedure to be performed, and may be used either in contact or out of contact with the target tissue.

The final delivery system is the air-cooled fiberoptic delivery system. This type of delivery system is unique to the erbium family of lasers. A conventional fiberoptic delivery system cannot transmit the wavelength of the erbium family of lasers, owing to the specific characteristics of the erbium wavelength. These special air-cooled fibers terminate in a handpiece with quartz or sapphire tips. These tips are used slightly (1-2mm) out of contact with the target tissue.

Since the introduction of the dLase 300, general practitioners have seen the number of wavelengths and manufacturers available to them increase by leaps and bounds. It has increased from one manufacturer of one wavelength to eight different manufacturers offering six different wavelengths7.

CLASSIFICATION OF LASERS

Lasers are divided into a number of classes depending upon the power or energy of the beam and the wavelength of the emitted radiation. Laser classification is based on the laser’s potential for causing immediate injury to the eye or skin and/or potential for causing fires from direct exposure to the beam or from reflections from diffuse reflective surfaces. Since August 1, 1976, commercially produced lasers have been classified and identified by labels affixed to the laser.

Radiation safety should be consulted on the appropriate laser classification and labeling. Lasers are classified using physical parameters of the laser, power, wavelength, and exposure duration.

- CLASS 1 LASERS

A Class 1 laser is considered safe based upon current medical knowledge. Under normal operating conditions Class 1 lasers or laser systems cannot produce a hazard and no safety measures are required. A Class 1 laser could also be a higher-class laser that is completely enclosed to prevent personnel exposure to the laser beam.

Example: Laser printers and CD players.

- CLASS 2 LASERS

A Class 2 laser or laser system is defined as operating in the visible region (400-700nm). These lasers are not inherently safe but protection against eye damage will normally be afforded by aversion responses including the blink reflex. Momentary viewing (exposure of 0.25 second or less) is not considered hazardous. Intentional extended viewing, however, is considered hazardous.

Example: Laser pointers, surveying lasers.

- CLASS 2A LASERS

Are special-purpose lasers not intended for viewing. Their power output is less than

1 mW. This class of lasers causes injury only when viewed directly for more than 1,000 seconds. The 1,000 seconds is spread over an 8-hour day, not continuous exposure.

Example: Many bar-code readers fall into this category.

- CLASS 3 LASERS

Class 3 lasers are medium-power lasers that pose a modest potential for injury. Class 3 laser users may be required to follow specific safety precautions and may require the wearing of safety equipment such as laser protective eye wear. Skin hazards normally do not exist for incidental exposures.

- CLASS 3A

A Class 3A laser emits higher levels of radiation and requires more stringent precautions than those necessary for Class 2 laser products. They differ from Class 2 laser products in that they emit more power in a beam of larger cross-section, so that when the output is viewed directly, the power of the beam entering the eye does not exceed that of a Class 2 laser product.

However, if the beam is viewed through larger diameter collecting optics (eg. binoculars) then the hazard is usually increased. For continuous wave (CW) output in the visible wavelength range, the output power from Class 3A lasers is limited to 5mW and the maximum irradiance (power density) is 25W.

Example: HeNe laser above 1 milliwatt but not exceeding 5 milliwatts radiant power, or some pocket laser pointers.

- CLASS 3B (RESTRICTED)

Class 3B (Restricted) lasers or laser systems operate at the same power levels as Class 3A but have higher levels (25 to 50Wm-2) of irradiance. They may be used in daylight conditions under the same controls as for Class 3A laser products. Where used in conditions of less illuminance (generally less than 10 lux), the appropriate safety controls are those specified for Class 3B laser products.

Example: Visible HeNe laser above 5 milliwatts but not exceeding 500 milliwatts radiant power.

- CLASS 3B

Class 3B lasers can emit either invisible or visible radiation and direct viewing is hazardous to the eye. Class 3B lasers are capable of causing eye injury either because their output is invisible and therefore aversion responses are not activated, or because the beam power is such that damage is done in a time shorter than the blink reflex (0.25s). Higher power lasers in this class may also cause skin burns. However, with laser wavelengths other than those in the ultraviolet region, the pain produced by rapid heating of the skin will usually evoke an aversion response sufficient to avoid such burns.

Class 3A is replaced by classes 1M and 2M. Class 3B represents maximal power output of 0.5 W. Examples include soft medical lasers (LLLT), laser light show equipment and laser measuring devices. A new classification is the class 3R which may also include some low level medical devices and targeting tissues but generally lasers of low power output than 3A.

- CLASS 4 LASERS

Class 4 lasers are high-power lasers with power levels greater than 500 mW radiant power, that pose a serious potential for injury of the eye and skin and require that users follow specific safety precautions and wear laser protective eyewear. They may operate in any part of the spectrum, as with the 3B.

Class 4 laser systems can produce a hazard not only from direct or specular reflections, but also from a diffuse reflection. A fire risk may also be associated with the use of such high powered systems and their use requires extreme caution.

Example: Most Nd:YAG Lasers.

Of these classifications, only Class 4 lasers are incorporated into systems used for material processing such as cutting, heat treatment, surfacing and welding. A Class 4 laser may be part of a system designed in such a manner as to be considered a Class 1 laser system. Such a system cannot under normal operating conditions, produce a hazard. This can be achieved using engineering controls such as enclosures, interlocks, and other mechanisms.

It is the duty of the manufacturer of the laser to classify the product and consequently to place warning labels on the product and to realise safety features such as key-locks and interlock connectors. However, if the user manipulates the laser product so that the class is changed, the user becomes responsible for a re-classification8.

LASER PHYSICS

The basic units, or quanta of light are called photons. From one perspective, the photons behave like tiny wavelets similar to sound wave pulses. A quantum of light can be depicted as an electromagnetic wave with an electric field oscillating up and down in the plane of the page and a magnetic field that moves in and out of the page.

Quanta of electromagnetic energy are classified as cosmic rays, gamma rays, light, microwaves, or radio waves. Light may be ultraviolet, visible (violet, blue, green, yellow, orange, red), or infrared. The physical property that determines this classification of electromagnetic energy is wavelength, which is the distance a photon travels while the electric field completes one complete oscillation. Cosmic rays have extremely short wavelengths (10-12m), whereas radio waves have very long wavelengths (103m). With current technology, laser light covers the range from about 0.1 to 10µm.

LIGHT ABSORPTION AND EMISSION

When light encounters matter, it can be deflected (reflected or scattered) or absorbed. If a photon is absorbed, its energy is not destroyed but rather used to increase the energy level of the absorbing atom or molecule. This idea is central to both laser physics and laser-tissue interactions.

An atom can absorb a photon which then ceases to exist, and a electron (e) within the atom jumps to a higher energy level (e*). This atom is thus pumped up to an excited state from the resting ground state. In the excited state, the atom is unstable and will soon spontaneously decay back to the ground state, releasing the stored energy in the form of an emitted photon. This process is called spontaneous emission.

The interval between absorption and re-emission is usually very short and defines the fluorescence lifetime of the atom. The figure below shows that the spontaneously emitted photon has less energy (a longer wavelength) than the absorbed photon. The difference in energy is usually turned into heat.

In any given atom, only certain orbits (energy levels) are allowed. When a photon is absorbed, the atom jumps to one of the allowable energy levels, which means that each type of atom or molecule can absorb only photons of exactly the right energy (or

wavelength). The result is that each species of atom or molecule has a unique absorption spectrum3.

LIGHT AMPLIFICATION BY STIMULATED EMISSION OF RADIATION

The process of lasing occurs when an excited atom can be stimulated to emit a photon before the process occurs spontaneously. When a photon of exactly the right energy (wavelength) enters the electromagnetic field of an excited atom, the incident photon triggers the decay of the excited electron to the lower energy state. This is accompanied by the release of the stored energy in the form a second photon. The first photon is not absorbed but continues to encounter another excited atom.

Stimulated emission can only occur when the incident photon has exactly the same energy as the released photon. Thus, the result of stimulated emission is two photons of identical wavelength travelling in the same direction. The release of the second photon is time – locked to the oscillations of the first photon, so that the two photons oscillate together in phase.

If a collection of atoms includes more photons that are pumped up into the excited state than remain in the resting state, a population inversion exists. This is a necessary condition for lasing. Now, the spontaneous emission of a photon by one atom will stimulate the release of a second photon in a second atom, and these two photons will trigger the release of two more photons ; these four then yield eight, eight yield sixteen, and so on. In a small space at the speed of light, this photon chain reaction produces a brief, intense flash of monochromatic (same wavelength) and coherent (same phase) light.

To contain and amplify the photon chain reaction that results from stimulated emission in a population of excited atoms, it is necessary to place this reaction within an optical cavity. An optical cavity consists of two parallel mirrors placed on either side of the laser medium. In this configuration, photons bounce off the mirrors and re-enter the medium to stimulate the release of more photons. If some form of energy is provided to continuously pump atoms up to the excited state, the population inversion can be maintained and high-intensity light circulating back and forth between the two mirrors can be generated.

The mirrors collimate the light, that is, photons exactly perpendicular to the mirrors re-enter the active medium, while those off axis leave the lasing process. Because the process is not 100 % efficient and some energy is converted into heat, it is necessary to provide some form of cooling.

If one mirror is totally reflective (M2) and the other mirror partially transmissive (M1), the light that escapes through M1 becomes the laser beam. If our circulating power is 1000 W and M1 is 10 % transmissive, then the power returning from M1 is 900 W, and the exciting beam is 100 W. A round trip through the lasing medium brings the power back up to 1000 W.

The active medium contains the homogeneous population of atoms or molecules that are pumped up to the excited state and are stimulated to lase. The exact species of atom or molecule determines the wavelength of the output beam. The active medium is suspended in the optical cavity as a gas, a liquid, or distributed in a solid state (eg : Crystal). Thus, stimulated emission within an optical cavity generates a collimated, coherent, and monochromatic beam of light. The laser is named for the contents of the active medium and their state of suspension, eg : CO2 gas laser or argon-ion gas laser.

The projection of the beam on the target is called the spot. A cross section of the beam is called the beam profile, a common way of describing spot geometry & the diameter of the spot is called the spot size. The size and shape of the lens determines the focal length and the spot size at the focal point. Beyond the focal point, the beam diverges and the power density or intensity decreases.

The terms focused and defocused when working with lasers refers to the position of the focal point in relation to the tissue plane.

When working on tissue, the laser should always be used either with the focal point positioned at the tissue surface (in focus) or positioned above the tissue surface (defocused or out of focus). The laser should never be positioned with the focal spot deep or within the tissue (prefocused) as this can lead to deeper thermal damage and undesirable tissue effects3.

Electromagnetic Spectrum And Wavelengths Of Lasers

## CHARACTERISTICS OF LASER

## Laser light is monochromatic, directional, and coherent

## The light emitted from a laser is monochromatic, that is, it is of one wavelength (color).  In contrast, ordinary white light is a combination of many different wavelengths (colors).

## Lasers emit light that is highly directional.  Laser light is emitted as a relatively narrow beam in a specific direction.  Ordinary light, such as coming from the sun, a light bulb, or a candle, is emitted in many directions away from the source.

## The light from a laser is said to be coherent, which means the wavelengths of the laser light are in phase in space and time.

## These three properties of laser light are what make it more of a hazard than ordinary light9.

## Lasers are also characterized by the duration of laser emission - continuous wave or pulsed laser. A Q-Switched laser is a pulsed laser which contains a shutter-like device that does not allow emission of laser light until opened. Energy is built-up in a Q-Switched laser and released by opening the device to produce a single, intense laser pulse.

|  | - CONTINUOUS WAVE (CW) : Lasers operate with a stable average beam power. In most higher power systems, one is able to adjust the power. In low power gas lasers, such as HeNe, the power level is fixed by the design, and performance usually degrades with long term use. |
| --- | --- |

- SINGLE PULSED : (normal mode) lasers generally have pulse durations of a few hundred microseconds to a few milliseconds. This mode of operation is sometimes referred to as long pulse or normal mode.

|  | - SINGLE PULSED Q-SWITCHED : lasers are the result of an intracavity delay (Q-switch cell) which allows the laser media to store a maximum of potential energy. Under optimum gain conditions, emission occurs in single pulses; typically of 10(-8) second time domain. These pulses will have high peak powers often in the range from 10(6) to 10(9) Watts peak. - REPETITIVELY PULSED : or scanning lasers generally involve the operation of pulsed laser performance operating at a fixed (or variable) pulse rates which may range from a few pulses per second to as high as 20,000 pulses per second. The direction of a CW laser can be scanned rapidly using optical scanning systems to produce the equivalent of a repetitively pulsed output at a given location. - MODE LOCKED : Lasers operate as a result of the resonant modes of an optical cavity which can effect the characteristics of the output beam. When the phases of different frequency modes are synchronized, i.e., "locked together", the different modes will interfere with one another to generate a beat effect. The result is a laser output which is observed as regularly spaced pulsations. Lasers operating in this mode-locked fashion, usually produce a train of regularly spaced pulses, each having a duration of 10(-15) (femto) to 10(-12) (pico) sec. A mode-locked laser can deliver extremely high peak powers than the same laser operating in the Q-switched mode. These pulses will have enormous peak powers often in the range from 10(12) Watts peak9. |
| --- | --- |

LASER PHYSICS

There are many types of lasers available for research, medical, industrial, and commercial uses.  Lasers are often described by the kind of lasing medium they use - solid state, gas, excimer, dye, or semiconductor.

|  | - SOLID STATE LASERS – Have lasing material distributed in a solid matrix, e.g., the ruby or neodymium-YAG (yttrium aluminum garnet) lasers. The neodymium-YAG laser emits infrared light at 1.064 micrometers. |
| --- | --- |
|  | - GAS LASERS – Helium and Helium-neon(HeNe), are the most common gas lasers which have a primary output of a visible red light. CO2 lasers emit energy in the far-infrared, 10.6 micrometers, used for cutting hard materials. |
|  | - EXCIMER LASERS – (Derived from the terms excited and dimers) Use reactive gases such as chlorine and fluorine mixed with inert gases such as argon, krypton, or xenon. When electrically stimulated, a pseudomolecule or dimer is produced and when lased, produces light in the ultraviolet range. |
|  | - DYE LASERS – Use complex organic dyes like rhodamine 6G in liquid solution or a suspension as lasing media. They are tuneable over a broad range of wavelengths. |
|  | - SEMICONDUCTOR LASERS – Sometimes called diode lasers, are not solid-state lasers. These electronic devices are generally very small and use low power. They may be built into larger arrays, e.g., the writing source in some laser printers or compact disk players9. |

LASER TYPE		         		          WAVELENGTH           COLOUR

Excimer lasers 	Argon Fluoride (ArF) 		193 nm 	Ultraviolet

Xenon Chloride (XeCl) 		308 nm 	Ultraviolet

Gas lasers 		Argon 				488 nm 	Blue

514 nm 	Blue-green

Helium Neon (HeNe) 		637 nm 	Red

Carbon Dioxide (CO2) 	      10,600 nm 	Infrared

Diode lasers 		InGaAsP				655 nm 	Red

GaAlAs			   670–830 nm	Red-infrared

GaAs					840 nm 	Infrared

InGaAs				980 nm 	Infrared

Solid state lasers 	Frequency-doubled 			337 nm 	Ultraviolet

Alexandrite

Potassium Titanyl			532 nm 	Green

Phosphate (KTP)

Neodymium:YAG 		        1,064 nm 	Infrared

Holmium:YAG		        2,100 nm 	Infrared

Erbium, chromium:YSGG 	        2,780 nm 	Infrared

Erbium:YSGG 		        2,790 nm 	Infrared

Erbium:YAG 		        2,940 nm 	Infrared

InGaAsP - Indium Gallium Arsenide Phosphorus, GaAlAs - Gallium Aluminum Arsenide, GaAs - Gallium Arsenide, InGaAs - Indium Gallium Arsenide

LASER INTERACTION WITH BIOLOGIC TISSUES

The oral cavity is a unique and complex environment, where hard and soft tissues exist in close proximity, within bacteria-laden saliva. All oral tissues are receptive to laser treatment, but the biophysics governing  laser-tissue interaction demands a knowledge of all factors involved in delivery of this modality. Through this knowledge, correct and appropriate treatment can be delivered in a predictable manner10.

The biologic factors that influence laser-tissue interactions are more extensive. Included among these are the optical properties of various tissue elements that govern how specific molecular and chemical components in tissue react with light energy. The optical properties of tissue elements determine the nature and extent of the tissue response through the processes of absorption, transmission, reflection, and scattering of the laser beam (Dederich, 1991).

Other factors involve the various types of physiologic and mechanical processes that occur as a result of energy transformations within tissue mass, the inflammatory response of tissues to noxious stimuli, tissue vascularity, and repair mechanisms. The incident light energy interacts with a medium (eg : oral tissue) that is denser than air, in one of the four ways as listed below3.

- TRANSMISSION

The inverse of absorption, wherein laser energy passes directly through the tissue with no interaction between the incident beam and the medium. The beam will emerge distally, unchanged or partially refracted. This effect is highly dependent on the wavelength of laser light.

- SCATTER

There is some interaction , but this is insufficient to cause complete attenuation of the beam. Scatter will cause some diminution of light energy with distance, together with a distortion in the beam, whereby rays proceed in an uncontrolled direction through the medium. Back-scatter of the laser beam can occur as it hits the tissue ; this is seen most in short wavelengths, eg : Diode, Nd : YAG ( > 50 % back-scatter ).

- REFLECTION

The beam redirects itself off the surface, producing no effect on the target tissue. The density of the medium, or angle of incidence being less than the refractive angle, results in a total reflection of the beam. In true reflection, the incident and emergence angles will be the same or, if the medium interface is rough or non-homogenous, some scatter may occur, as seen in caries detection, which uses reflected light to measure the degree of sound tooth structure.

- ABSORPTION

The amount of energy absorbed by a tissue depends on its characteristics, such as pigmentation and water content, laser wavelength and emission mode. The incident energy of the beam is attenuated by the medium and transferred into another form. Haemoglobin reflects red wavelengths, imparting colour to the arterial blood. It therefore is strongly absorbed by blue and green wavelengths.

Venous blood, containing less oxygen, absorbs more red light and appears darker. The pigment melanin, which imparts colour to the skin, is strongly absorbed by short wavelengths, and water, the universally present molecule, has varying degrees of absorption by different wavelengths.

Dental tissues have different amounts of water content by weight. A ranking from lowest to highest would show enamel (with 2% to 3%), dentin, bone, calculus, caries, and soft tissue (about 70%). In general, the shorter wavelengths (from about 500-1000nm) are readily absorbed in pigmented tissue and blood elements10.

CURRENT AND POTENTIAL APPLICATIONS OF LASERS IN DENTISTRY 11

LASER TYPE 		   CURRENT/POTENTIAL DENTAL APPLICATION

EXCIMER LASERS

Argon Fluoride (ArF)		Hard tissue ablation, Dental calculus removal

Xenon Chloride (XeCl)

GAS LASERS

Argon (Ar)                                    	Curing of composite materials, Tooth whitening, Intraoral soft tissue surgery,Sulcular debridement (subgingival curettage in periodontitis and peri-implantitis)

Helium Neon (HeNe) 	Analgesia, Treatment of dentin hypersensitivity, Aphthous ulcer treatment

Carbon Dioxide (CO2)	Intraoral and implant soft tissue surgery, Aphthous ulcer treatment, Removal

of gingival melanin pigmentation, Treatment of dentin hypersensitivity, Analgesia

DIODE LASERS

Indium Gallium Arsenide			Caries and calculus detection

Phosphorus (InGaAsP)

Galium Aluminum				Intraoral general and implant soft tissue

Arsenide (GaAlAs) and 			surgery, Sulcular debridement

Galium Arsenide (GaAs) (subgingival curettage in periodontitis and peri-implantitis), Analgesia,

Treatment of dentin hypersensitivity, Pulpotomy, Root canal disinfection,

Aphthous ulcer treatment, Removal of gingival melanin pigmentation

SOLID STATE LASERS

Frequency-doubled				Selective ablation of dental plaque and calculus

Alexandrite

Neodymium:YAG	Intraoral soft tissue surgery, Sulcular

(Nd:YAG)	debridement (subgingival curettage in periodontitis), Analgesia, Treatment of dentin hypersensitivity, Pulpotomy, Root canal disinfection, Removal of enamel caries, Aphthous ulcer treatment, Removal of gingival melanin pigmentation

ERBIUM GROUP	Caries removal and cavitypreparation,

Erbium:YAG (Er:YAG),			Modification of enamel and dentin

Erbium:YSGG (Er:YSGG),	surfaces, Intraoral general and implant soft

Erbium,chromium:YSGG	tissue surgery, Sulcular debridement

(Er,Cr:YSGG) 	(subgingival curettage in periodontitis and peri-implantitis), Scaling of root surfaces, Osseous surgery, Treatment of dentin hypersensitivity, Analgesia, Pulpotomy, Root canal treatment and

disinfection, Aphthous ulcer treatment, Removal of gingival melanin/metal-tattoo pigmentation

APPLICATION OF LASER FOR PERIODONTAL TREATMENT

Scaling and root planning is the traditional method of controlling subgingival microflora for management of periodontal diseases. The objectives of subgingival debridement are to eliminate not only the adherent and unattached bacterial plaque, but also deposits of calculus. However, removal of calculus using conventional hand instruments has been reported to be incomplete and rather time consuming. In order to improve the effectiveness and efficiency of root surface debridement, various devices such as sonic and ultrasonic scalers, and more recently lasers have been used. Many studies have demonstrated that sonic and ultrasonic instrumentation, when compared with manual instrumentation, results in equal and superior treatment outcomes.

In periodontal pockets, the root surfaces are contaminated with an accumulation of plaque and calculus, as well as infiltration of bacteria and bacterial endotoxins into cementum. Complete removal of these harmful substances is essential for the healing of periodontal tissue. Formation of biofilms on the exposed root surface within periodontal pockets impedes the infiltration of antibiotics, and therefore mechanical disruption of the biofilm is necessary during periodontal treatment.

The use of lasers for periodontal treatment becomes more complicated because the periodontium consists of both hard and soft tissues.

Among the many lasers available, high power lasers such as CO2, Nd:YAG and diode lasers can be used in periodontics because of their excellent soft tissue ablation and hemostatic characteristics. However, when they are applied to the root surface or alveolar bone, carbonization and thermal damage have been reported. Therefore the use of these lasers is limited to gingivectomy, frenectomy and similar soft tissue procedures including the removal of melanin pigmentation of gingiva12.

Recently, Er:YAG and Erbium-Chromium doped:Yittrium-Selenium-Gallium-Garnet (Er,Cr:YSGG) laser scaling was introduced as an alternative or an adjunctive to conventional scaling and root debridement (Figure 2 and 3). Of all the lasers available, the absorption of the Er:YAG and Er,Cr:YSGG lasers in water is nearly the highest. These lasers effectively ablate all biologic tissues that contain water molecules. The erbium laser group has emerged as a promising laser system for periodontal indications.

Several clinical studies have reported the application of Er:YAG laser for periodontal treatment. Watanabe et al. demonstrated efficient calculus removal with no side effects and uneventful reduction of pocket after Er:YAG scaling. Schwarz et al. reported that equal or slightly better results were observed at six months after laser treatment of periodontal pockets, compared to conventional mechanical debridement using hand scalers and found significantly higher reduction of bleeding on probing scores and improvements in clinical attachment level after laser treatment.

Schwarz et al. also demonstrated that nonsurgical periodontal treatment with laser alone and a combination of Er:YAG laser and scaling and root planning using hand instruments may result in clinically and statistically significant improvements in the clinical parameters with no difference between two treatments, 12 months after treatment.

Er:YAG laser was also proposed for the implant maintenance, taking advantage of its bactericidal or decontamination effect. Peri-implant infection results in inflammation of the surrounding soft tissues and can induce a breakdown of the implant supporting alveolar bone. It is associated with the presence of a subgingival microflora, which seems to be quite similar to that in periodontal pockets and contains a large variety of Gram-negative anaerobic bacteria.

Matsuyama et al. performed debridement of implant abutment surface by Er:YAG laser and reported effective removal of plaque and calculus without producing damage to the implant surface. Also, Kreisler et al. observed a non excessive heat generation on the implant surfaces and effective decontamination by means of the Er:YAG laser.

Even though successful experimental results and clinical results have been reported so far with the Er:YAG laser, further studies are required to better understand the effects on periodontium for its safe and effective application during the periodontal treatment.

Therefore, randomised controlled clinical trials and more basic studies have to be encouraged and performed to determine the most optimal and safest parameters for laser treatment12.

ADVANTAGES AND DISADVANTAGES

Advantages of laser treatment in periodontics are effective and efficient soft and hard tissue ablation with a greater hemostasis, bactericidal effect, minimal wound contraction, minimal collateral damages with reduced use of local analgesia. In addition, the small popping sound of the lasers in action with Er:YAG seems to produce less stress to patients than the high pitch vibration sound of most of the ultrasonic devices.

Despite numerous advantages of using lasers, the use of laser also has disadvantages that require precautions to be taken during clinical application. Laser irradiation can interact with tissues even in the non-contact mode, which means that laser beams may reach the patient's eye and other tissues surrounding the target in the oral cavity. Clinicians should be careful to prevent inadvertent irradiation to these tissues, especially to the eyes. Protective eyewear specific for the wavelength of the laser in use must be worn by patient, operator, and assistant.

It is recommended that dental laser users to attend certification courses provided by some dental laser organisations and follow laser safety guidelines such as the Laser Code of Practice from the Hong Kong Surgical Laser Association. A good understanding in laser wavelength characteristics, tissue interaction and laser device specification provide a platform for achieving the best results12.

THE SURGICAL TECHNIQUE

Dental lasers fall into two basic categories :

- Those that work solely in the noncontact mode, either focused or defocused, such as CO2 lasers and

- Those that work either in the contact or noncontact modes, the contact mode being focused and the noncontact mode being defocused. This group includes the lasers delivered by  fiberoptics such as argon, Ho:YAG, and Nd:YAG.

- NONCONTACT LASERS

This includes carbon dioxide lasers which can be divided into those CO2 lasers that work by articulated arm technology, those that operate via hollow waveguide technology, and those that operate by hand-held technology.

Most CO2 lasers applicable in medicine can also be used effectively in dentistry. Carbon dioxide lasers work primarily in a noncontact mode. They can also be used either in a focused or defocused mode.

FOCUSED MODE :  Is when the laser beam hits the tissue at its focal point or the smallest diameter. This diameter is dependent on the size of the lens used. Most CO2 laser systems today have a lens that can focus the beam down to spot sizes ranging

from 0.1mm to 0.35mm or larger. This focused mode can also be referred to as a cut mode. For example, the cut mode is used when performing biopsies.

DEFOCUSED MODE : By defocusing the laser beam or moving the focal spot away from the tissue plane, the beam size that hits the tissue has a greater diameter, thus causing a wider area of tissue to be vaporized. By defocusing the beam, however, laser intensity or power density is reduced. A defocused beam, for example, is used when performing frenectomies, or for the removal of inflammatory papillary hyperplasia3.

Because all optical fibers have a high absorptive index for CO2 laser light, true fiberoptic delivery cannot be used. Instead, some CO2 lasers use articulating arms with properly aligned mirrors at knuckle joints and then a handpiece with a specific lens is used to create the desired spot size in the focused mode. Because of their sometimes cumbersome nature, CO2 laser systems that use articulated arms are often difficult to use within the oral cavity. With articulated arm systems and those that use hand-held technology, the focal distance to tissue can vary from 1 to 3 cm.

- CONTACT VS NONCONTACT MODES

Noncontact surgery relies completely on the pigments and water present in the tissue that are responsible for the absorption model of each wavelength. Directing sufficient energy densities to the target tissue will then create the desired tissue effect. The creation and appearance of char on the tissue surface will also increase the energy absorption at the surface and will improve the efficiency. Directing the energy at one spot will gradually heat an increasing mass of tissue, and moving the energy along a path with single or multiple strokes will create an incision. Painting a large area with the energy will contour and sculpt the tissue. Decreasing the energy will allow the operator to coagulate and create hemostasis in most surgical sites in the mouth.

Noncontact surgery requires significantly greater power settings than contact surgery for the fiberoptic delivered lasers to achieve the equivalent tissue effect. The fiber tip must be well cleaved and kept free of debris. With proper technique, the results are the same as those for contact surgery ; the techniques and principles are similar, only the feel is different. During noncontact surgery, the clinician operates with visual control with the aid of an aiming beam or by observing the tissue effect being created.

In contact mode the treatment probe is held firmly and perpendicular to the tissue surface with the tip pressed into the tissue surface. The depth to which treatment probe is pressed depends upon the depth and the site of the target tissue. This minimises attenuation of beams and reduces perfusion3.

CARBON DIOXIDE LASER

- CHARACTERISTICS

The CO2 laser is a gas-active medium laser that incorporates a sealed tube containing a gaseous mixture with CO2 molecules pumped via electrical discharge current. The light energy, whose wavelength is 10,600 nm, is placed at the end of the mid-infrared invisible nonionizing portion of the spectrum, and it is delivered through a hollow tube-like waveguide in continuous or gated pulsed mode.

This wavelength is well absorbed by water, second only to the erbium family. It can easily cut and coagulate soft tissue, with a shallow depth of penetration, which is important when treating mucosal lesions. Because this wavelength was one of the earliest used in general medical surgery, there are numerous published papers verifying its efficacy.

The CO2 laser cannot be delivered in a conventional optic fiber. The North American products use a hollow waveguide with a handpiece and accessory tips. The laser energy is conducted through  the waveguide and is focused onto the surgical site in a noncontact fashion. The loss of tactile sensation could pose a disadvantage for the surgeon, but the tissue ablation can be precise with careful technique.

Large lesions can be treated using a simple back and forth motion ; the procedure proceeds quickly because there is no need to touch the tissue. The noncontact mode thus has an advantage when treating movable oral structures, such as the tongue and floor of the mouth. After the surgery is completed, many clinicians use a defocused beam to place a biologic bandage called an eschar on the wound surface.

This wavelength has the highest absorption in hydroxyapatite, compared to any dental laser, and is about 1000 times greater than erbium. Therefore, tooth structure adjacent to a soft-tissue surgical site must be shielded from the incident laser beam ; usually a metal instrument placed in the sulcus provides the protection. The continuous wave emission and delivery system technology of CO2 devices limit hard-tissue applications because carbonization and crazing of tooth structure can occur due to the long pulse duration and low peak powers.

However, ongoing research using experimental devices with extremely short pulses show favourable results for surface modification and strengthening of tooth enamel for increased caries resistance.

The CO2 laser is absorbed at the tissue surface with very little scatter or penetration. Since ablation is mainly due to the action of heat generation, carbonization easily occurs on the irradiated surface, but the heat produced does not scatter.

Therefore, the CO2 laser produces a relatively thin layer of thermally changed tissue (coagulation) around the ablated site.

The width of the coagulated layer was reported to be 100–300 µm in an incision of pigskin with the continuous mode CO2 laser at 6 W13. Tissue penetration from this laser irradiation will be approximately 0.5 mm deep, depending on power density2. In the field of periodontics, there have been several reports on laser application for gingivectomy and gingivoplasty since late 1980s14,15,16,17.

Transmission of the CO2 laser through optical fibers was very difficult and therefore the CO2 laser system previously employed mirror systems using articulated arms for laser beam delivery. Recently, new flexible fiber optic delivery and hollow tube wave-guiding systems have been developed, along with the development of contact tips. These advances may render the use of the CO2 laser for periodontal pockets possible in the near future.

BASIC AND CLINICAL STUDIES ON ROOT SURFACE

PREPARATION AND CALCULUS REMOVAL

Several basic studies have shown the effects of continuous wave CO2 laser irradiation on root surfaces. The continuous wave CO2 laser readily produces carbonization, melting, and cracking of root cementum and dentin 18,19,20,21,22.

Scanning electron micrograph of the root surface after CO2 laser irradiation (continuous wave, focused beam, 0.5 W), showing numerous typical microcracks.

Tucker et al.23 in 1996 evaluated the effects of the CO2 laser on calculus in vitro and reported that the pulsed CO2 laser at 6 W and 20 Hz (pulse duration:0.01 s) was able to remove dental plaque on the root surface, whereas only melting and carbonization occurred on the dental calculus of extracted teeth.

Misra et al.24 in 1999 examined the root conditioning effects of the defocused mode CO2 laser after scaling and root planing in vitro. Laser irradiation at 3 W for 1 s completely removed the smear layer with minimal change in the diameter of the dentinal tubules; however, irradiation times of 1.2 and 1.4s produced surface charring and carbonization, and were totally ineffective in exposing the dentinal tubules.

Barone et al.25 in 2002 investigated the effects of the pulsed defocus mode CO2 laser. The CO2 laser at 2.0 W and 4 Hz with 4.0 mm spot size did not result in any extensive damages to the root surface, which was flat and smooth with apparent fusion of the smear layer. They concluded that the pulsed defocus mode may present the advantage of decontaminating the root surface.

Crepsi et al.26 in 2002 reported that after the pulsed defocus mode CO2 laser treatment at 2 W and 1 Hz, the periodontally diseased root surface showed the highest number of tightly attached fibroblasts compared with the nontreated control and scaling and root planing (SRP) alone. They concluded that pulsed defocus mode CO2 laser treatment combined with mechanical instrumentation constitutes a useful tool for root conditioning.

Miyazaki et al.27 in 2003 applied CO2 laser irradiation for pocket treatment on the external surface of the marginal gingiva.

They used a continuous wave mode CO2 laser (2.0 W, 120 s) and reported decreased inflammation and probing depth after treatment.

However, so far there have been no reported clinical studies on application of the CO2 laser in periodontal pockets.

Nd:YAG LASER

- CHARACTERISTICS

The Nd:YAG laser is a free-running pulsed wave laser with a wavelength of 1,064 nm. Unlike the CO2 and Er:YAG lasers, the Nd:YAG laser has low absorption in water, and the energy scatters or penetrates into the biological tissues. In water, the Nd:YAG laser will theoretically penetrate to a depth of 60 mm before it is attenuated to 10% of its original strength.

The photothermal effect of the Nd:YAG laser is useful for soft tissue surgery. Due to the characteristics of penetration and thermogenesis, the Nd:YAG laser produces a relatively thick coagulation layer on the lased soft tissue surface, and thereby shows strong hemostasis. Hence, the Nd:YAG laser is basically effective for ablation of potentially hemorrhagic soft tissue. The width of the coagulation layer was 0.3–0.8 mm in an incision of bovine oral soft tissue in vitro at 3–10 W.

In dentistry, soft tissue surgery using the Nd:YAG laser has been widely accepted 28,29,30,31,32. In 1990, the FDA approved soft tissue removal by means of a pulsed Nd:YAG laser 33. White et al. in 1991 successfully used the Nd:YAG laser for intraoral soft tissue application without anesthesia, and with minimal bleeding compared to scalpel surgery.

It is easy to deliver the Nd:YAG laser by a flexible optical fiber with a contact tip of 400 µm (core diameter: 320 µm) suitable for pocket insertion. Basic research and clinical trials have been performed on periodontal pocket curettage and root surface debridement. In 1997, the FDA approved sulcular debridement by means of a pulsed Nd:YAG laser.

The Nd:YAG laser is not suitable for ablation of intact hard tissues. However, caries removal using this laser is possible to some extent 34,35. White et al.35 reported a safe and effective procedure for selective removal of enamel caries with the Nd:YAG laser, and the FDA approved removal of enamel (first degree) caries using an Nd:YAG laser in 199936. As the Nd:YAG laser is well absorbed by dark substances, Indian ink or other kinds of black pigment are often applied to increase the efficiency of ablation.

BASIC STUDIES

Calculus removal:

In an in vitro study, Tseng & Liew37,38,39 demonstrated that partial removal and detachment of the calculus from the root surface was achieved with the Nd:YAG laser at 2.0 or 2.75 W and 20 Hz. However, melting of calculus and thermal damage was noted in localized areas of the original cementum and even dentin after irradiation at

high power37. They also reported that, after laser irradiation, removal of remaining calculus by curettes was facilitated37,38.

Arcoria & Vitasek-Arcoria40 in 1992 assessed the effects of the Nd:YAG laser on calculus removal from root surfaces in vitro. Nd:YAG laser irradiation was performed at the calculus–cementum interface in contact mode at 1.5 or 3.0 W (15 Hz) with the tip inclined 450 to the root surface. The integrity of the calculus–root surface attachment was not appreciably affected by 1.5 W irradiation, whereas 3.0 W irradiation detached the calculus without root surface damage, similarly to conventional hand instrumentation.

In an in vitro study, Radvar et al.41 in 1995 reported that Nd:YAG laser irradiation, perpendicular to the root surface, at 0.5 – 2.0 W (50 or 100 mJ/pulse, 10 or 20 Hz) caused greater ablation of calculus than either cementum or dentin. However, the specimens with complete evaporation of calculus also showed some degree of damage to the underlying cementum.

As subgingival calculus is originally dark in color, the Nd:YAG laser has the advantage of being absorbed well into subgingival calculus. Use of higher energy levels may ablate calculus more efficiently but may be inappropriate for clinical usage due to increased thermal side-effects.

Root surface treatment :

Several basic studies have shown the effects of Nd:YAG laser irradiation on root surfaces. Nd:YAG laser produced thermal changes such as carbonization, melting, and resolidification of root cementum and dentin 42,43,44,45

Spencer et al.45 in 1992 reported a decrease in the protein/mineral ratio in cementum samples lased with the Nd:YAG laser at 0.8 W (80 mJ/pulse and 10 Hz) using FTIR analysis. They considered that the decreased protein/mineral ratio and the potential surface contamination with protein by-products might ultimately affect cell reattachment at the cementum surface. Later, they suggested that the protein by-products would be cyanamide and cyanate ions46,47.

Thomas et al.48 in 1994 also reported that the nondiseased root surfaces treated at 1.5 W (75 mJ/ pulse, 20 Hz) with water coolant exhibited a significantly decreased fibroblast attachment, but the root planing or polishing with air-powder abrasive that followed, increased the cell attachment. They concluded that alterations in the laser-irradiated surface are reversible and additional root treatment following laser irradiation appears essential to render the root surface biocompatible. FTIR spectroscopy of the lased root surface revealed reduction of the Amide II band, suggesting denaturation of surface protein upon laser exposure.

Nd:YAG LASER TREATMENT HAS BEEN REPORTED TO REMOVE THE SMEAR LAYER ON THE ROOT SURFACE

Wilder-Smith et al.49 in 1995 examined the effect of smear layer removal using the Nd:YAG laser at 5 W (pulse durations and intervals of 0.1s;5 Hz and calculated energy output 1,000 mJ/pulse) without coolant. Although the smear layer was removed without microstructural changes of the hard tissue, a significant rise in the intrapulpal and root surface temperature occurred. They concluded that the irradiation parameters used may not be appropriate for clinical use.

Interestingly, Fukuda et al.50 in 1994 performed Nd:YAG laser irradiation to the periodontally diseased root surface of extracted teeth, and reported that the Nd:YAG laser at 0.3–0.5 W (30–50 mJ/pulse, 10 Hz, 2–4 s) could inactivate the endotoxin in the superficial layer of the root surface.

Liu et al.51 in 2002 explored the in vitro effectiveness of the Nd:YAG laser for the elimination of cementum-bound endotoxin by measuring interleukin (IL)-1b changes in stimulated monocytes. Nd:YAG laser varying between 50 mJ/ pulse and 10 Hz and 150 mJ/pulse and 20 Hz for 2 min did not seem to be effective in destroying diseased cementum endotoxin. However, in their study, the thermal effect of Nd:YAG laser would have been weakened by noncontact irradiation with a long distance (1.5 cm) to the moistened cementum particles.

Clinical studies

Cobb et al.52  performed Nd:YAG laser irradiation into periodontal pockets in vivo in 18 teeth from eight patients, and examined the effects of root debridement using a power of 1.75–3.0 W (87.5–150 mJ/ pulse, 20 Hz), either alone or in combination with manual instrumentation. Application of the Nd:YAG laser resulted in ineffective and patchy removal of calculus from the root surface. The remaining calculus had a characteristic porous surface due to carbonization, melting, and resolidification but the surface was free of microbial plaque. They suggested that laser therapy should be followed by mechanical debridement.

They reported the effect of Nd:YAG laser on both root surface and subgingival microflora in periodontal pockets, employing teeth that required extraction. DNA probe analysis showed a decreased number of bacteria in periodontal pockets after laser treatment. In addition, there were root surface alterations similar to those observed in the in vitro studies.

Gold & Vilardi53 performed Nd:YAG laser curettage in vivo, and studied microscopically 24 specimens of gingival tissue from six patients following the laser application at 1.25 and 1.75 W (62.5 and 87.5 mJ/ pulse, 20 Hz). The pulsed Nd:YAG laser removed pocket-lining epithelium in moderately deep periodontal pockets without causing necrosis or carbonization of the underlying connective tissue.

Neil & Mellonig54 in 1997 performed a double-blind, randomized, controlled clinical trial for sulcular debridement on 186 teeth from 10 patients using a split-mouth design to compare the adjunctive use of the pulsed Nd:YAG laser at 2 W (80 mJ/pulse, 25 Hz) to scaling and root planing alone in the nonsurgical treatment of moderate to advanced adult periodontitis.

The reduction of probing depth for 6 months after treatment was similar between the scaling and root planing plus laser therapy and scaling and root planing alone, but the scaling and root planing plus laser therapy showed significantly greater improvements in gingival index and gingival bleeding index at specific time points. The mean attachment level of laser-treated pockets also showed a tendency to improve steadily for 6 months, whereas the scaling and root planing group showed a reduction in attachment level.

Miyazaki et al.55 in 2003 compared the effectiveness of the Nd:YAG and CO2 laser treatment to that of ultrasonic scaling in periodontal pockets of chronic periodontitis patients. The 41 sites of 18 patients were randomly assigned for treatment with the Nd:YAG laser alone (2.0 W: 100 mJ/pulse and 20 Hz, 120 s), the CO2 laser alone (2.0 W, 120 s), or ultrasonic scaling alone (maximum power, 120 s). Decreased inflammation and probing depth were observed in all three groups after treatment. There were significant decreases in both Porphyromonas gingivalis and the amount of gingival crevicular fluid in the Nd:YAG and scaling groups.

The Nd:YAG group also tended to show a decrease in IL-1 level. Nd:YAG laser and ultrasonic scaling treatments showed significant improvements regarding the clinical parameters and subgingival microflora compared to the baseline, but no significant difference was observed between the three groups.

Various combinations of irradiation parameters have been reported for the clinical use of the Nd:YAG laser for pocket treatment56. White56 in 2002 recommended 1.5 W (100 mJ/pulse, 15 Hz) irradiation for removal of the sulcular diseased tissue and 2.0 W (100 mJ/pulse, 20 Hz) irradiation for coagulation of soft tissue wall after mechanical debridement.

Coluzzi57 in 2002 recommended laser soft tissue curettage at 1.8 W (30 mJ/pulse, 60 Hz) after mechanical debridement, followed by irradiation at 2 W (100 mJ/pulse, 20 Hz) for hemostasis and bacterial reduction. Gutknecht et al.58 suggested the use of the Nd:YAG laser at 2 W (100 mJ/pulse, 20 Hz) for curettage before mechanical debridement to reduce the risk of bacteremia after the scaling and root planing procedures, and to facilitate mechanical debridement.

Scanning electron micrograph of the root surface after Nd:YAG laser irradiation (1 W: 100 mJ/pulse and 10 Hz), showing typical melting and resolidification with a porous structure.

Er:YAG LASER

- Characteristics

The Er:YAG laser was introduced in 1974 by Zharikov et al. as a solid-state laser that generates a light with a wavelength of 2,940 nm. Of all lasers emitting in the near- and mid-infrared spectral range, the absorption of the Er:YAG laser in water is the greatest because its 2,940 nm wavelength coincides with the large absorption band for water.

The absorption coefficient of water of the Er:YAG laser is theoretically 10,000 and 15,000–20,000 times higher than that of the CO2 and the Nd:YAG lasers, respectively 59,60. Also, as part of the apatite component, OH groups show a relatively high absorption at 2,940 nm, although the maximum absorption is around 2,800 nm61. Since the Er:YAG laser is well absorbed by all biological tissues that contain water molecules, this laser is indicated not only for the treatment of soft tissues but also for ablation of hard tissues.

In dentistry, the free-running pulsed Er:YAG laser has already been used clinically for caries removal and cavity preparation62,63,64,65,66 and soft tissue treatment 67,68. The FDA approved the pulsed Er:YAG laser for hard tissue treatment such as caries removal and cavity preparation in 199769, unchanged for soft tissue surgery and sulcular debridement in 1999 and for osseous surgery in 2004.

The high absorption of the Er:YAG laser into water minimizes thermal influences on the surrounding tissues during irradiation. When the Er:YAG laser was used for an incision of pigskin in a noncontact mode, it showed formation of a thermally changed layer of only 10–50 µm70. In the case of hard tissue procedures, some degree of heat generation is inevitable with the Er:YAG laser, since the Er:YAG laser emits in the infrared region and hard tissues have very low water content.

However, the use of water coolant minimizes heat generation by cooling the irradiated area and absorbing excessive laser energy71,72,73. In addition, a water spray facilitates hard tissue ablation by keeping the target moist. Er:YAG laser irradiation using water irrigation has been reported to produce an altered layer 5–15 µm in width on cementum and dentin surfaces.

The recently introduced Erbium,Chromium-doped:Yttrium-Scandium-Gallium-Garnet (Er,Cr:YSGG) laser with 2,780 nm wavelength and the Erbium-doped:Yttrium-Scandium-Gallium-Garnet (Er:YSGG) laser with 2,790 nm wavelength, which are more highly absorbed by OH ions than water molecules74, are expected to have a performance similar to that of the Er:YAG laser.

MECHANISM OF TISSUE ABLATION WITH THE Er:YAG LASER

A mechanism of biological tissue ablation with the Er:YAG laser has been proposed, based on the optical properties of its emission wavelength and morphologic features of the surface ablated by Er:YAG laser. During Er:YAG laser irradiation, the laser energy is absorbed selectively by water molecules and hydrous organic components of biological tissues, causing evaporation of water and organic components, resulting in thermal effects due to the heat generated by this process (‘photothermal evaporation’). Moreover, in hard tissue procedures, the water vapor production induces an increase of internal pressure within the tissue, resulting in explosive expansion called ‘microexplosion’75,76.

These dynamic effects cause mechanical tissue collapse, resulting in a ‘thermomechanical’ or ‘photomechanical’ ablation77. This phenomenon has also been referred to as ‘water-mediated explosive ablation’78.

Unlike CO2 laser (10,600 nm) irradiation, the Er:YAG laser is capable of ablating hard tissues. Fried et al.78 in 2000 suggested that the difference between the Er:YAG laser and CO2 laser (10,600 nm) in the mechanism of enamel ablation is based primarily on the principal absorber in the hard tissue. He reported that primary absorption in water results in water-mediated ablation, and primary absorption in the bulk of enamel rods results in melting and vaporization.

The absorption of the Er:YAG laser by inorganic components (hydroxyapatite) is much lower than that of the CO2 laser77. Thus, in the hard tissue ablation with the Er:YAG laser, the absorption into water and hydrous organic components rapidly occurs before heat accumulation caused by absorption into inorganic components takes place, resulting in thermo-mechanical, explosive ablation. The excellent ablation effect of the Er:YAG laser of both soft and hard tissues has received a lot of attention in the field of periodontal therapy, and has been extensively researched.

Photograph of the root surface after single pulse irradiation of the Er:YAG laser at various energy levels. Single pulse Er:YAG laser irradiation was performed under water irrigation in a contact mode using a conventional cylindrical tip with a 600 µm diameter on the intact root surface at an energy density of 3.5, 7.1, 10.6, 17.7, 28.2, and 42.4 J/cm2 per pulse (energy output of 10, 20, 30, 50, 80, and 120 mJ/pulse, respectively) starting from the upper left in the figure. The ablation defect became deeper according to the increase in energy output. The ablation defect shows chalky appearance without carbonization.

Scanning electron micrograph of the root surface after single pulse Er:YAG laser irradiation. The irradiation was performed in a contact mode at an energy density of 42.4 J/cm2 (energy output of 120 mJ/pulse) under water irrigation. The lased site shows crater-like ablation defect with a well defined edge and microstructured bottom, and without major crackings and thermal alterations.

Schematic view of hard tissue ablation by Er:YAG laser irradiation. During Er:YAG laser irradiation, the laser energy is absorbed selectively by water molecules and hydrous organic components of the biological hard tissues and causes evaporation of water (photothermal evaporation). Further, in the case of hard tissue, the vapour produces an increase in the internal pressure leading to micro-explosions and consequent mechanical tissue collapse (thermomechanical or photomechanical ablation).

The Er:YAG laser irradiation produces a very thin altered layer on the ablated surface, which consists of two distinct sub-layers: a superficial, significantly altered layer and a deeper, less affected (intermediate) layer. In the directly irradiated, superficial layer, micro-cracking, disorganization and slight recrystallization of the original apatites and reduction of surrounding organic matrix are evident after microexplosion, while the intermediate layer, which receives less energy, suffers mainly from the effects of energy accumulation, such as heat and subsequent microexplosion. The deep layer under the intermediate layer shows no change. W: water molecules. BA : biological apatites. PM : protein matrix.

Before irradiation

Superficial layer

Intermediate layer

Deep layer

After irradiation

Er:YAG laser irradiation

Basic studies

Removal of subgingival calculus: Dental calculus contains water in its structural micropores as well as in its intrinsic components. Since the Er:YAG laser has the ability to ablate dental hard tissues such as enamel and dentin, the laser was expected to be capable of removing dental calculus at much lower energy levels.

In 1994, Aoki et al.79 first documented the ability of the Er:YAG laser to remove subgingival calculus in vitro. They showed that the pulsed Er:YAG laser used with water irrigation was capable of removing subgingival calculus from the root surface effectively at 30 mJ/pulse (energy density of single pulse at the tip: 10.6 J/cm2 per pulse) and 10 Hz, in the contact mode, directed perpendicular to the root surface using a conventional cylindrical contact tip 600 µm in diameter.

In addition, ablation of the tooth substance following laser scaling was generally observed within cementum, with a slight rise in temperature of pulpal side during scaling. Their study suggested the potential for clinical application of the Er:YAG laser in subgingival scaling.

Schwarz et al.80 in 2003 compared the degree of calculus removal with in vivo Er:YAG laser irradiation at 160 mJ/pulse (energy output 120 mJ/pulse and chisel tip size 1.65 x 0.5 mm; calculated energy density 14.5 J/cm2 per pulse) and 10 Hz with removal after water spray or scaling and root planing with hand instruments.

The study used roots of teeth that had been planned for extraction due to severe periodontal destruction. The laser treatment was performed in a coronal to apical direction in parallel paths, with the fiber tip inclined 15–200 to the root surface. The Er:YAG laser treatment provided selective subgingival calculus removal to a level equivalent to that provided by scaling and root planing.

However, Frentzen et al.81 in 2002 reported that, although Er:YAG laser scaling achieved complete debridement clinically, laser scaling at a panel setting of 160 mJ/pulse (output energy 100 or 120 mJ/pulse and calculated energy density 18.8 or 14.5 J/cm2 per pulse in the use of chisel tip 1.1 x 0.5 mm or 1.65 x 0.5 mm, respectively) and 10 Hz with water spray resulted in an increased loss of cementum and dentin in vitro compared to mechanical scaling.

They considered that this loss should be taken into account in the clinical situation. The crater depth of the treated root surface was approximately 40 and 80 µm with the use of tip 1.65 and 1.10, respectively.

Regarding their results, Ishikawa82 in pointed out that Frentzen et al. employed a relatively high energy output for Er:YAG laser scaling, and commented that although the efficiency of laser scaling can be easily improved by using a higher output power, caution should be used when deciding on the power output, considering a balance of effectiveness and unnecessary tissue removal.

Improvement of the effectiveness of laser scaling should rely on other variables, such as pulse repetition rate and pulse duration, rather than only on an increase of energy output.

Contact tips used for Er:YAG laser irradiation. a) Conventional cylindrical contact tip, which has a curvature of 800, a round end with a diameter of 600 µm and approximately 65% laser transmission (800 Tip – 600 Micron, Er:YAG laser DELightTM; HOYA ConBio, CA, USA, and Er:YAG laser DentlightTM; HOYA Photonics Corp.,Tokyo, Japan). b) Conventional cylindrical contact tip with a tapered end which enables lateral irradiation, used for soft tissue incision and periodontal pocket curettage (S600T, Er:YAG laser Erwin AdvErLTM; Morita Mfg. Corp., Kyoto, Japan). c) Chisel-type tip which has a rectangular end with dimensions 1.65 x 0.5 mm, and is generally used with a contra-angle handpiece for pocket treatment (Truncated-core end 1.65 x 0.5 mm and periodontal handpiece, KEY LaserTM; KaVo, Biberach, Germany).

ROOT SURFACE ALTERATION FOLLOWING ER:YAG LASER IRRADIATION

The Er:YAG laser does not cause carbonization of the irradiated root surface, but it has been demonstrated that the ablated surface becomes chalky after drying due to micro-irregularities on the lased surface83,84. The surface of Er:YAG laser-treated calculus under water coolant has been reported to exhibit a micro-irregular appearance without melting and carbonization, likely due to the effects of the mechanical ablation 83.

Israel et al.84 in 1997 reported that the root surface treated with the Er:YAG laser at 2.5–15 mJ/pulse and 20 Hz with water spray in the noncontact irradiation exhibited an etched appearance with exposure of numerous tufts and/or fiber bundles of mineralized collagen as well as sharply defined microfractures of mineralized structure.

Fujii et al.85 in 1998 showed a microstructured root surface with denaturation of collagen fibers up to a depth of 15µm in cementum, following single-pulse, perpendicular contact Er:YAG laser irradiation at 75 mJ/pulse (calculated energy density: 26.5 J/cm2 per pulse) under water spray.

Photographs of the root surfaces before and after Er:YAG laser scaling of subgingival calculus. Er:YAG laser scaling was performed in vitro at an energy density of 10.6 J/cm2 per pulse (energy output of 60 mJ/pulse) and 30 Hz under water spray. The laser was applied using a chisel-type tip (1.4 x 0.45 mm) in contact mode, maintaining the tip oblique to the root surface at an angle of approximately 200, and moving the tip in a back-and forth motion. Macroscopically, the laser-scaled area appears smooth and slightly chalky, but without any major thermal changes such as carbonization. a) Before laser scaling. b) After laser scaling. Part of the subgingival calculus was intentionally left on the edge of the treated root surface.

PIONEER WORKS WITH ERBIUM:YAG LASER

Hibst et al. was the first to report the use of the Er:YAG laser in 198886 for ablation of dental hard tissues. During the same period, several researchers examined the effectiveness of this laser for hard tissue ablation. In 1989, Paghdiwala prepared holes in enamel and dentin with low energy and without water cooling, which showed that the Er:YAG laser was capable of ablating dental hard tissues87. Kayano et al. suggested that the Er:YAG laser could be used for cavity preparation88,89.

In the same year, Keller and Hibst demonstrated in extracted teeth that the enamel and dentin removal could be performed effectively with minimal thermal damage to the adjacent tissues. Several studies showed that the Er:YAG laser is effective and safe for caries removal. In 1995, Kumazaki reported effective caries removal by Er:YAG laser90. In a scanning electron microscopic (SEM) study, Ishikawa et al. observed effective removal of softened carious dentine by Er:YAG laser ablation, with no smear layer on the irradiated area91.

Pelagalli et al. also demonstrated by SEM that the resultant surface morphology and ability to remove caries were similar after treatment with Er:YAG laser and mechanical drill92. Furthermore, other studies reported effective caries removal by this laser, with minimal thermal damage to the adjacent dentin.

GINGIVAL MELANIN AND GINGIVAL DISCOLORATION REMOVAL

The normal colour of gingiva is pink. Gingival hyperpigmentation is mostly caused by the physiologic deposition of melanin by melanocytes. In patients with a ‘’gummy smile’’, this pigmentation causes an aesthetic problem. Methods to remove this pigmentation may vary, but it seems that the most reliable & satisfactory procedure is by laser ablation. The Er:YAG laser is capable of excellent soft tissue ablation, which makes it suitable for this kind of pigmentation removal93.

In a study done by Azzeh et al in 2007, six white patients complaining of dark-brown to black gingival hyperpigmentation were included in the study. Laser ablation was performed by an Er:YAG laser, with a setting of 250 mJ and 15 Hz with water and air in a defocused mode without topical or local anaesthesia. In all patients, no pain, discomfort and bleeding complications were found and no recurrence was observed during the follow-up period which ranged from 6 to 18 months94. Clinically results indicated safe and effective melanin pigmentation ablation95,96.

Furthermore, the Er:YAG laser was utilized for removal of abnormal gingival discoloration, namely metal tattoos. It is suspected that dental treatment procedures or restorative materials cause this kind of discoloration.

A recent report indicated that this kind of discolored gingiva contained silver sulfate, tin sulfate, and pieces of iron, as observed by electron probe microscopy97. Apparently, the iron must have originated from cutting steel instruments, since it is not a component of silver alloy and without surgical removal of the gingival tissues containing the fragments, the discoloration will still remain85. Effective removal of discolored gingiva together with metal fragments was reported possible using an Er:YAG laser, with no pain or gingival recession98.

Removal of gingival melanin hyperpigmentation using Er:YAG laser. (a) Gingival tissue with a band of pigmentation before surgery. (b) Er:YAG laser irradiation at 27 mJ/pulse and 20–30 Hz with water spray under topical anesthesia for removal of melanin hyperpigmentation. (c) Immediately after surgery. Effective removal of pigmented gingival tissue without major thermal damages such as carbonization and coagulation. (d) Favorable healing of the treated site at 1 week after the procedure.

BONE TISSUE APPLICATION

The application of the Er:YAG laser on bone in oral and periodontal surgery is not very common. There are few reports regarding the use of this laser for bone ablation. Nelson et al. reported that the Er:YAG laser ablated bone effectively with minimal thermal damage to the adjacent tissues99. The ability of bone tissue removal with minimal chemical and morphological changes to the irradiated and surrounding surfaces was demonstrated previously.

A typical irregular pattern, which consisted of biological apatites surrounded by organic matrix, was observed in the irradiated bone and this may have aided in uneventful healing100. Sasaki et al. showed, histologically, a thin altered layer produced by Er:YAG laser on the irradiated rat calvaria bone surface. The irregular bone morphology after irradiation and absence of toxic substances may have promoted the adhesion of plasma proteins during the initial stages of healing. On the other hand, the bone ablated using CO2 laser showed extensive thermal effects100.

In another study from our research group, Pourzarandian et al. demonstrated that the initial bone healing following Er:YAG laser irradiation occurred faster than that after mechanical bur drilling and CO2 laser irradiation in rats, observed by light and transmission electron microscopy.

This laser system has been demonstrated to be useful for bone ablation and osseous recontouring during periodontal surgery. There is a possibility that the bone is biostimulated after Er:YAG laser irradiation, but further experiments have to be carried out to elucidate the exact mechanism of Er:YAG laser irradiation in bone tissue.

GRANULATION TISSUE REMOVAL

In 1995, Williams et al. used CO2 laser for removal of granulation tissue and connective tissue from interproximal craters101. Sasaki et al. suggested the possibility of granulation tissue removal by Er:YAG laser during periodontal flap surgery102. Although it may be a promising tool for granulation tissue removal in periodontal pocket sites, there appears to be no clinical study using Er:YAG laser. Further research is required to show efficacy of granulation tissue removal without any thermal damage to the adjacent surrounding tissue including alveolar bone and root surfaces.

IMPLANT MAINTENANCE

The Er:YAG laser was also proposed for implant maintenance, taking advantage of its bactericidal effect, technical simplicity and absence of postoperative pain and edema103,104. Peri-implant infection results in inflammation of the surrounding soft tissues and can induce a breakdown of the implant-supporting bone. It is associated with the presence of a subgingival microflora, which is quite similar to that in periodontal pockets and contains a large variety of Gram negative anaerobic bacteria.

Kreisler et al. demonstrated in vitro the high bactericidal potential of Er:YAG laser on titanium implants with different surface characteristics105. Matsuyama et al. performed debridement of implant abutment surface by this laser and reported effective removal of plaque and calculus without injury to the surface106. Also, Kreisler et al. in 2002 observed nonexcessive heat generation on the implant surfaces and effective decontamination by means of the Er:YAG laser.

In vivo, Schwarz et al. demonstrated the effectiveness of this laser treatment to remove subgingival calculus from surfaces of implant fixtures without any thermal damage107.

Bone surgery using Er:YAG laser. (a) Exostosis before resection. (b) Exostosis removal by Er:YAG laser irradiation at 100 mJ/pulse and 10 Hz with saline water spray. (c) Site immediately after effective removal of exostosis.

Granulation tissue debridement using Er:YAG laser with water irrigation during flap operation. (a) Presence of granulation tissues in vertical bone defects. (b) Effective granulation tissue removal from vertical bone defects after Er:YAG laser debridement at 100 mJ/pulse and 10 Hz with saline water spray.

Histological observation by light microscopy of bone surface after laser irradiation. (a) Histological photomicrograph after Er:YAG laser irradiation at 100 mJ/pulse and 10 Hz (1 W) under water irrigation. Groove-like appearance and thin changed layer of irradiated bone surface. (b) Histological photomicrograph after CO2 laser irradiation at 1 W. Extensive thermal effect produced on bone surface. A dark carbonized layer is evident. From Sasaki et al. Ultrastructural analysis of bone tissue irradiated by Er:YAG laser.

Removal of abnormal gingival discoloration, namely metal tattoos, using an Er:YAG laser in combination with a surgical microscope. (a) Metal tattoos at marginal area of central incisors before the treatment. (b) Treatment site immediately after irradiation at 40 mJ/pulse and 30 Hz with water spray under local anesthesia. (c) Four weeks after removal. Effective removal of discolored gingiva without any recession.

DIFFERENT CHARACTERISTICS OF ERBIUM AND NEODYMIUM:YAG LASERS

ERBIUM:YAG LASER	           NEODYMIUM:YAG LASER

Wavelength 		2940 nm infrared 			1060 nm near infrared

Body interaction 	Surface absorption type 		Deep penetration type

Ablation type 		Photo-mechanical ablation 		Photo-thermal ablation

(Disintegration by heat		(Ablation by heat burning)

expanded distortion of water)

Characteristics	Very thin surface interaction 		The laser penetrates into

water to a depth of 60 mm

Absorbance rate is about  		Scattering in soft tissues

10 times higher than that

of CO2 laser

Less temperature rise with 		Causing ‘cooking effect’

water irrigation

Less tissue degeneration 	Vaporizing and removing tissue

Hard tissue preparation 		Selective absorption in dark

substance

Advantage 	Applicable to both hard and 		Best suited for  primary                        soft tissues     				coagulative properties                 No carbonization 			Selective and fine cutting

with contact probes

Easy to perform scaling and 		Optically transmittable fiber

caries treatment

Sterilization effect

Smooth laser knife

Optically transmittable fiber

Disadvantage 		Less hemostasis 			Difficulty in judging

the depth of penetration

Impossible to prepare cavities in teeth

DIODE LASERS

- Characteristics

The diode laser is a solid-state semiconductor laser that typically uses a combination of Gallium (Ga), Arsenide (Ar), and other elements such as Aluminum (Al) and Indium (In) to change electrical energy into light energy. The wavelength range is about 800–980 nm. The laser is emitted in continuous-wave and gated-pulsed modes, and is usually operated in a contact method using a flexible fiber optic delivery system. Laser light at 800–980 nm is poorly absorbed in water, but highly absorbed in hemoglobin and other pigments108.

Since the diode basically does not interact with dental hard tissues, the laser is an excellent soft tissue surgical laser, indicated for cutting and coagulating gingiva and oral mucosa, for soft tissue curettage or sulcular debridement. The FDA approved oral soft tissue surgery in 1995 and sulcular debridement in 1998 by means of a diode laser (GaAlAs 810 nm).

The diode laser exhibits thermal effects using the ‘hot-tip’ effect caused by heat accumulation at the end of the fiber, and produces a relatively thick coagulation layer on the treated surface108. The usage is quite similar to electrocauterization.

Tissue penetration of a diode laser is less than that of the Nd:YAG laser, while the rate of heat generation is higher, resulting in deeper coagulation and more charring on the surface compared to the Nd:YAG laser. The width of the coagulation layer was reported to be in excess of 1.0 mm in an incision of bovine oral soft tissue in vitro. The advantages of diode lasers are the smaller size of the units as well as the lower financial costs.

Basic studies

Concerning the effects on the root surface, Kreisler et al.109 in 2001 examined the periodontal ligament cell attachment to the 810 nm diode laser-treated root surface. After scaling and root planing the periodontally diseased root surface with curettes followed by air-powder abrasive treatment, the laser group received diode laser irradiation at 1 W in the continuous wave mode for 20 s and the control group was left unirradiated. There was no significant difference between the laser and the control groups in cell attachment. This finding suggested that the diode laser did not produce any deleterious effect on the root surface. Thus, it is generally considered that diode laser surgery can be performed safely in close proximity to dental hard tissue.

However, Kreisler et al.110 in 2002 further evaluated possible morphologic alterations of the root surfaces with a human blood film after noncontact Ga-AlAs-diode laser (809 nm) irradiation (0.5–2.5 W, continuous wave, 10–30 s).

Interestingly, they reported that lasing dry or saline-moistened root specimens resulted in no detectable alterations ; however, the blood-coated specimens showed severe damage depending on the irradiation conditions.

Irradiation at 1 W and below had barely any negative effect on the root surface, whereas irradiation at 1.5 W and higher resulted in partial or total carbonization. Kreisler et al. also examined intra-pulpal temperature elevations during diode laser (809 nm GaAlAs) irradiation on the root surface, performing laser irradiation at 0.5–2.5 W in the continuous wave mode for 120 s. Temperature elevations between 0.5 and 32.00C were registered in an energy and time-dependent manner. They reported the risk of temperature elevation of the pulpal side during diode laser irradiation on the root surface.

Schwarz et al.111 in 2003 performed in vivo GaAlAs diode laser treatment (810 nm, 1.8 W, pulsed, pulse/pause relation 1 : 10) on periodontally diseased roots of teeth considered for extraction due to severe periodontal destruction and examined the degree of calculus removal after extraction.

They reported that diode laser was unsuitable for calculus removal and altered the root surface in an undesirable manner. Thus, diode lasers at a high energy level, especially

in a continuous mode, can cause root surface alterations in the presence of blood and elevated temperatures, depending on the power employed.

DIODE LASER, BASIC AND CLINICAL STUDIES ON ROOT SURFACE AND PERIODONTAL POCKET TREATMENTS

Clinical studies

Some studies have demonstrated that a diode laser facilitated bacterial elimination from periodontal pockets, resulting in better healing. Moritz et al.112 reported pocket irradiation with a diode laser (805 nm) following scaling. Irradiation with the diode laser at a power output of 2.5 W in pulsed mode (50 Hz, pulse duration 10 ms), produced considerable bacterial elimination from periodontal pockets at a much higher level than the scaling alone group, especially in terms of A. actinomycetemcomitans.

Moritz et al. in 1998113 also performed a clinical study using a diode laser (805 nm) as an adjunctive treatment for periodontal pockets in order to reduce or eliminate bacteria. The pulsed irradiation at 2.5 W (50 Hz, pulse duration 10 ms) was performed three times, at 1 week, 2 months, and 4 months interval after scaling, while the control group was rinsed with H2O2. After 6 months, bacterial reduction in the laser therapy group was significantly higher than in the control group. The improvement of bleeding on probing scores and pocket depths were greater in the laser group. They concluded that diode laser therapy, in combination with scaling, supports healing of periodontal pockets by eliminating bacteria.

Regarding clinical use of the diode laser for pocket treatment, Coluzzi114 recommended laser soft tissue curettage at 0.4 W in continuous wave mode after mechanical debridement of root surface, followed by irradiation at 0.6 W for hemostasis and bacterial reduction, while Gutknecht et al. in 2001 suggested the use of a diode laser at 2 W in continuous wavemode for curettage before mechanical debridement. Further detailed studies are required to establish proper irradiation conditions, including the use of water coolant, for periodontal pocket therapy with diode lasers.

Other applications

Hibst et al.115,116 developed a novel laser device for caries detection (Diagnodent, KaVo, Biberach, Germany), which uses laser fluorescence induced by the 655 nm InGaAsP diode laser. It was suggested that caries-associated bacteria or their byproducts might be the source of reaction to the increasing fluorescence.

Hibst et al.116 in 2001 identified the source of red excited fluorescence present in caries lesions as porphyrins, especially proto-porphyrin IX, which are products of oral bacteria, such as Prevotella intermedia and P. gingivalis. They also suggested another application of laser fluorescence for calculus detection. Recently, the 655 nm diode laser system has been reported to be useful for detection of calculus that includes a significant amount of bacteria or their byproducts.

Krause et al.117 observed that laser fluorescence values decreased significantly after in vitro scaling of extracted teeth, and the values were strongly correlated with the presence of calculus. Schwarz et al.80 reported that the Er:YAG laser, combined with a calculus detection system with fluorescence induced by 655 nm InGaAsP diode laser, provided selective subgingival calculus removal to a level equivalent to that provided by scaling and root planing.

Traditionally, calculus detection has been performed manually by judging the ruggedness of the root surface using a periodontal probe. The laser fluorescence probe may be a novel, valuable tool for clinical detection of calculus in the near future.

ARGON LASER

- Characteristics

The argon laser uses argon ion gas as an active medium and is fiber optically delivered in continuous wave and gated pulsed modes. This laser has two wavelengths, 488 nm (blue) and 514 nm (blue-green), in the spectrum of visible light. The argon laser is poorly absorbed in water and therefore does not interact with dental hard tissues. However, it is well absorbed in pigmented tissues, including haemoglobin and melanin, and in pigmented bacteria.

Although not widely used in periodontal therapy, in operative dentistry the 488 nm argon laser is commonly used for curing composite resin and bleaching teeth in the dental office. It has also been studied in the application for caries prevention. The argon laser was approved by the FDA for oral soft tissue surgery and curing of composite materials in 1991 and for tooth whitening in 1995118.

Basic and clinical studies

Henry et al.119,120 reported that the argon laser at a low level has bactericidal effect on Prevotella and Porphyromonas species in the presence of oxygen. They suggested that low doses of argon laser radiation may be effective in the treatment of clinical infections caused by biofilm-associated species of Prevotella and Porphyromonas.

In a clinical study, Finkbeiner121 treated a total of 1,328 pockets from 30 patients, with argon laser pocket thermolysis in combination with mechanical root planing. Argon laser treatment was performed using a 300-µm fiber in contact at 0.4 W, for 20–30 s per pocket, with coaxial irrigation. He reported that the 4–5 mm pockets were reduced by a mean of 1.62 mm, the 6–7 mm pockets by 2.85 mm, and the 8–9 mm pockets by 3.30 mm.

Considering the advantages of eradication of pigmented bacteria, this laser may be useful in the treatment of periodontal pockets. Further in vitro and in vivo studies are required to demonstrate the clinical benefits of this laser.

ALEXANDRITE LASER

- Characteristics

The Alexandrite laser is a solid-state laser employing a gemstone called Alexandrite, which is chromium doped:Beryllium-Aluminum-Oxide chrysoberyl (Cr+3;BeAl2O4) and is one of the few trichroic minerals.

In 1995, Rechmann & Henning first reported that the frequency-doubled Alexandrite laser (wavelength 337 nm, pulse duration 100 ns, double spikes, q-switched) could remove dental calculus in a completely selective mode without ablating the underlying enamel or cementum. Based on the difference in spectral region of fluorescence emission from dentin and that from subgingival calculus, they assumed that the wavelength of Alexandrite laser may be favorable for selective calculus ablation.

Their studies revealed that the Alexandrite laser at a fluence of 1 J/cm2 and pulse repetition rate of 55 Hz under water-cooling could selectively ablate supra and subgingival calculus as well as dental plaque. This laser has a wavelength in the ultraviolet spectrum and therefore does not produce any morphologic damage to enamel surface or root cementum, although extremely slight compositional change such as minimal reduction of amide II band is detected in the lased cementum by FTIR spectroscopy analysis.

Rechmann et al.122 in 2000 also demonstrated that there was no pulpal damage after removal of calculus with the Alexandrite laser at 1.5–6 J/cm2 and 70 Hz (pulse duration 1 ls) under water-cooling in dogs. However, the mechanism of selective ablation has not been clarified yet.

The development of this laser for clinical use is widely expected due to its excellent ability for selective calculus removal from the tooth or root surface without ablating the tooth structure. However, there is concern regarding use of light in the ultraviolet spectral region. Further studies are required to demonstrate the safety and effectiveness of this laser in clinical usage and to develop a laser apparatus which is appropriate for clinical use.

EXCIMER LASERS

- Characteristics

Excimer lasers are lasers that use a noble-gas halide, which is unstable, to generate radiation, usually in the ultraviolet region of the spectrum. Excimer laser wavelength depends on the chemical component, serving as the medium of the laser. It has been suggested that tissue ablation occurs in the nonthermal process of photoablation, likely due to an instantaneous increase of temperature or a straight combination of chemical elements.

In an in vitro experiment, Frentzen et al.123 demonstrated that the ArF excimer laser, wavelength 193 nm, could effectively remove dental calculus without causing any damage to the underlying surface. The cementum surface was clean, and only a slight roughness could be observed after irradiation, supporting the use of excimer lasers for laser scaling. Folwaczny et al.124 have reported that the 308 nm wavelength XeCl excimer laser could effectively ablate dental calculus without thermal damages or smear layer production.

Recently, flexible quartz glass fibres for XeCl excimer laser delivery systems have become available. However, apparatus cost and size still constitute an obstacle for clinical application of these lasers. Furthermore, ultraviolet rays should be used with caution, as they may have deleterious effects on biological tissues.

LASERS IN DENTAL IMPLANTOLOGY

The origin of implant dentistry can be traced back to China, 4000 years ago when bamboo was inserted into the jaw bone for fixed tooth replacements. In 1985, the publication of Tissue-Integrated Prostheses : Osseointegration in Clinical Dentistry by Branemark et al ushered in the era of osseointegration.

Parallelism in the expansion of implant and laser dentistry in clinical practice is apparent. As advocates for laser dentistry continue to seek new ways to use the technology and as more practitioners become involved in implant dentistry, it is logical to seek the concurrent use of both technologies in clinical practice.

The advantages of using lasers in implant dentistry are the same as for any other soft tissue dental procedure, which include increased hemostasis, minimal damage to the surrounding tissue, reduced swelling, reduced infection, and reduced post-operative pain. Due to the hemostasis provided by lasers, there is significant advantage of improved visibility during surgery. The increasing popularity of the erbium family of lasers, wtih their hard tissue ablation capability,has added the potential for its use for osteotomy and decontamination of infected and ailing implant bodies. However, there also is the potential for obliteration of the attached gingiva if this technology is

overused. It is important to maintain and preserve attached gingiva around implants whenever possible.

Surgical lasers can be used in a variety of ways with regard to implantology, ranging from placement, second stage recovery and gingival management, through to the treatment of peri-implantitis. Within this range of usage, dependant on wavelength employed, exists the ablation of target tissue and the ability to reduce bacterial contamination.

Titanium as a metal exhibits reflectivity to incident light energy. With regard to the wavelength of current lasers, the reflectivity is lowest in the range 780-900 nm, rising as the wavelength increases towards 10,600 nm (CO2 laser emission). This would suggest that shorter wavelengths are most damaging, as the low reflectivity would allow greater heat effects to build up, and is in keeping with studies carried out with the Nd:YAG laser.

However there is evidence to suggest that the diode wavelength group, delivered in low power CW values (1-2 Watts average power), cause minimal damage to the implant or surrounding bone. This is explained by the fact that the Nd:YAG, Er,Cr:YSGG and Er:YAG emission modes (FRP), result in high peak power values and heat production (> several hundred 0C).

Despite the damaging effects of carbon dioxide laser light on bone, several studies have borne out the high reflectivity of titanium to this wavelength in reporting low thermal effects on the metal surface and non-damaging effects on the metal composition.

Based on laser-tissue interaction characteristics, all laser wavelengths are suitable for the second stage recovery of implants, provided care is exercised to avoid contact with the implant body.

Suggested energy levels of one to two Watts (CW diode), 150 mJ/15 pps (Nd:YAG), 200-250 mJ/10 pps (erbium group and one to two Watts (CO2), appear to be appropriate in removing gingival tissue overlying the implant cover screw. The prime advantages of laser use in this procedure would be hemostasis, facilitating easier visual access to the cover screw, production of a protective coagulum as an aid to healing and patient comfort during and after treatment.

Minor surgical correction of the gingival margin can be carried out, to assist adequate implant exposure or to establish the correct emergence profile of the trans-mucosal element125.

LASER USE IN PERI-IMPLANTITIS

It is well known that adherent bacterial plaque and calculus develop on the surface of implant abutments, as in natural teeth. The maintenance treatment is required to keep the peri-implant tissue healthy in implant therapy. However, mechanical instruments such as metal curettes and ultrasonic scalers are prohibited for decontamination of titanium implant surfaces, since they easily damage the titanium surface.

Peri-implantitis is a condition associated with progressive peri-implant bone loss. Bacteria may penetrate the peri-implant tissues, and if the infection is left untreated, advanced bone loss may lead to implant failure. Many treatment methods have been recommended for peri-implant bony lesions, but solid evidence for successful management is scant: only 2 clinical studies on the surgical treatment of peri-implant bony defects involving larger groups of patients have been published.

The early stages of mucositis as well as peri-implantitis can be treated with different antimicrobial agents (e.g., chlorhexidine digluconate, citric acid). In more advanced alveolar defects, surgical treatment is also required.

Adequate removal of bacteria from the implant surfaces is a prerequisite for new bone formation.

Successful decontamination of implant surfaces by chemical or mechanical means (or both) enables some peri-implant bone regeneration to occur.

Evidence of the therapeutic effects of local application of tetracycline fibres around failing implants is inconclusive. Systemic administration of antibiotics may not be desirable or effective because of pharmacologic limitations (i.e., bacteria resistance, ineffective dosage, etc.). The use of curettes and ultrasonic instruments for decontamination has been criticized because they cause damage to the implant surface.

As for air-powder abrasive instruments, they should be used with extreme caution because patients face an increased risk of emphysema, especially when the instruments are used in the decontamination of deep alveolar bony defects. This treatment method may also damage the surface of hydroxyapatite coated implants. In recent years, lasers have been used to decontaminate implant surfaces. The essential points concerning the effectiveness and safety of different lasers used for this treatment are briefly outlined below.

SURFACE DECONTAMINATION USING LASER APPLICATIONS

Soft and hard lasers can be used to decontaminate implant surfaces.

Soft lasers: Significant antimicrobial effects have been demonstrated when peri-implant pockets are irrigated with toluidine blue followed by irradiation with a diode soft laser (905 nm for one minute). The toluidine blue sensitizes the bacterial cell membrane to laser light.

Hard lasers: The application of the contact Nd:YAG (neodymium:yttrium aluminium garnet) laser leads to adequate decontamination of the implant but may cause significant melting and crater-like formation of the implant surface. A significant temperature increase at the implant surface during Nd:YAG laser irradiation has been reported. Therefore, the application of contact Nd:YAG laser for peri-implant surgery is contraindicated.

The CO2 laser may be useful in the treatment of peri-implant lesions because of its bactericidal effect. With the CO2 laser, there is no significant increase in the surface temperature of the implant or any alteration to the implant surface pattern (as viewed using scanning electron microscopy). Access flap surgery and final implant surface laser irradiation should be used for implant decontamination immediately before bone augmentation procedures (Figs. 1–4).

Diode lasers (980 nm) do not damage titanium surfaces even when they are used in high-power settings as that of hard lasers. They can be used for the removal of peri-implant gingival overgrowths and to decontaminate implant surfaces before bony augmentations. The advantage of this type of laser is that it comes in small, reliable delivery units. The use of diode lasers with 810 nm wavelength and at a high-power setting may damage the implant surface, and for this reason such a laser must be used with special care to successfully treat peri-implantitis.

A bactericidal effect has been observed with the erbium:YAG laser, but some authors have reported damage to the implant surface after irradiation. Some negative effects on implant surfaces have also been reported after frequency-doubled Alexandrite-laser irradiation.

In conclusion, lasers play a valuable role in implant surface decontamination, a necessary step in the treatment of peri-implantitis. Knowledge of laser–tissue interaction is important to prevent complications126.

Figure 1: Deep infrabony peri-implant Defect. 	           Figure 2: Decontamination of the implant

surface  using  defocused beam of the  CO2 laser.

Figure 3: Augmentation of the lesion with 		Figure 4: Radiological examination 3 years

bone grafting material Bio-Oss (Osteohealth 		after surgery shows complete new bone fill

Co., Shirley, NY).				    	in the treated peri-implant lesion.

LOW LEVEL LASER THERAPY

- LLT TECHNOLOGY

The major components of a Low level laser therapy (LLLT) system are the laser device itself, a delivery system, and a controller. All common commercially available LLLT systems use semiconductor diode lasers.

These are generally variants of either Gallium:Aluminium:Arsenide (GaAlAs) which emit in the near infrared spectrum (wavelength 700-940 nm), or Indium : Gallium : Arsenide : Phosphorus (InGaAsP) devices which emit in the red portion of the visible spectrum range (wavelength 600-680 nm). Power outputs are typically 10-50 mW, when measured at the level of the diode laser itself, although the final useable output (from the handpiece) will be less because of losses in the delivery system. For this

reason, calibration of the laser system using an external power meter is an important quality assurance measure.

The output of diode lasers is controlled electronically, and in state-of-the-art equipment, microprocessor controllers are used. These allow the user to alter most aspects of the laser output other than wavelength (eg : power, pulse frequency, exposure time).

The diode laser device itself is quite compact, being approximately the size of an 8 mm diameter light emitting diode. The conversion efficiency (i.e, electrical energy to laser energy) is approximately 20 per cent, the remainder being dissipated as heat. Since the increased temperature of the laser device during operation alters both the output power (and to a lesser extent the wavelength), it is critical that the output of the laser diode is monitored so that the control circuitry can make the necessary adjustments to maintain a constant output. This is usually accomplished using a phototransistor which is fitted within the heatsink of the laser device. With an adequate heatsink and control system, the negative effect of temperature on laser output at the level of the treatment beam can virtually be eliminated.

The beam profile from a typical diode laser is rectangular (100x4μm) with a high divergence on the long axis (20 degrees from the centre axis), and a low divergence on the short axis (2 degrees). This gives a highly divergent ‘sweep’ profile. To obtain a more useful beam, a series of lenses or a self-focusing graded index fibre can be used in front of the device to either deliver the treatment beam itself or to direct the laser output into a small diameter flexible optical fibre or a solid light guide (such as a curing light tip). Customized devices have been designed to achieve a sweeping pattern of the treatment beam across the treatment area, although these are not used commonly.

Whatever the delivery system used, it is important that the components which come into direct contact with patients are able to be protected adequately with a laser-transmissive disposable barrier, or are themselves autoclavable or disposable. Similarly, it should be possible for the clinician to activate the laser into treatment mode without breaching asepsis.

Some units employ footswitches or light-operated switches to allow hands-free use. Laser units used for LLLT are generally classified as Class III or Class IIIb in terms of the optical hazard which they pose to staff and patients. Because a low power treatment beam can be focused by the eye to give a high power density on the retina, the optical hazard is sufficiently great, that laser safety standards (e.g., Australian Standards AS2211-1991 and AS4173-1994)† mandate the wearing of appropriate protective glasses by patients and clinicians during treatment. Glasses are available which provide protection against common LLLT wavelengths in both the visible and near infrared spectrum.

The use of LLLT in the treatment of dentinal hypersensitivity and periodontal ligament pain during orthodontic tooth movement has been shown in clinical trials to be both safe and effective. There is accumulating evidence which indicates the potential of lethal laser photosensitization as a technique for the destruction of cariogenic and other microorganisms within the mouth without causing undue thermal

stress to the tooth. Improvements in the design of LLLT equipment are necessary to enable these various techniques to be accomplished within an adequate time frame and without breaching crossinfection control requirements. Given the low - technology, low-cost characteristics of LLLT, the future for hard tissue LLLT applications is promising. As with soft tissue applications of LLLT, efforts should be directed toward investigating the precise dosimetry required for therapeutic laser effects, in order to achieve standardization of treatment protocols128.

LLLT IN DENTINAL HYPER SENSITIVITY

Dentinal hypersensitivity (DH) is one of the most common causes of dental pain. Low level laser therapy (LLLT) for the treatment of DH has been developed. This relies upon laser-induced changes to neural transmission networks within the dental pulp, rather than alterations to the exposed dentine surface, as is the case with other treatment modalities.

There have been several studies of the effect of LLLT on dentine hypersensitivity. Most studies have used GaAlAs laser treatment, and have demonstrated desensitization of hypersensitive cervical dentine, with an efficacy rate of approximately 90 per cent.

LLLT is efficacious when either the tooth crown or the root apex is irradiated. In a double blind study127, subjects with DH received LLLT involving both the apex and the cervical areas of the tooth. Treatment was repeated at one week, two-week and eight-week intervals. There was a reduction in both tactile and thermal sensitivity, and these were even more marked in the second and eighth weeks.

There is also evidence that LLLT delivered using HeNe lasers can reduce DH. In a randomized, double-blind clinical study involving 19 subjects, 8 hypersensitivity was

assessed using a mechanical stimulus (a sharp explorer), and a thermal stimulus (a blast of cold air from a dental syringe), while pulpal vitality was measured using an electrical stimulus.

Immediately following laser treatment and for three months thereafter, the subjects perceived level of discomfort decreased. HeNe laser treatment reduced dentine hypersensitivity to air by 63 percent and to mechanical stimulation by 61 per cent over three months. All teeth remained vital after laser treatment, with no adverse reactions or complications.

Based on a consideration of the optical properties of tooth structure, the phenomenon of ‘laser desensitization’ must be as a direct result from effects of LLLT on neural networks within dental pulp, rather than any accompanying thermal effects.

Since dentine and enamel are partly transparent to near infrared laser wavelengths, the majority of the deposited laser energy is transmitted through to the pulp, rather than dissipated at the tooth surface or at other optical boundaries. The mechanisms of the desensitizing effect at the cellular level have not yet been fully investigated; however, it is likely that inhibition of nociceptive signals arising from peripheral nerves is a major component of the therapeutic effect. Indeed, LLLT has been shown to block the depolarization of C-fibre afferents in dental pulp.

It has been suggested that LLLT may elicit descending inhibition in the central nervous system, in addition to local effects on nerve conduction, although this has yet to be proven128.

LLLT IN WOUND HEALING

Like photosynthesis, the correct wavelengths and power of light at certain intensities for an appropriate period of time can increase ATP production. Cell membrane perturbation could lead to permeability changes and second messenger activity resulting in functional changes such as increased syntheses, increased secretion and motility changes. Red and near infrared light seem to be the most ideal wavelengths. Red light and near infrared light acts on the mitochondria and at the cell membrane.

In in-vitro and animal LLLT wound healing studies comparing wavelengths, red consistently is more effective.

Shorter wavelengths are not as good and are more expensive to produce and have poor penetration; overall, they are a poor choice. Near infrared light, while not quite as good, do penetrate better than the red wavelengths and are available in higher powers and at low prices. According to live in-vivo experiments at Uniformed Services University Bethesda Maryland (a US military research centre) 810nm is the best penetrating wavelength. It also happens to work well in LLLT nerve regeneration studies they are doing.

Low-energy photon irradiation by light in the far-red to near-IR spectral range with low-energy (LLLT) lasers or LED arrays has been found to modulate various biological processes in cell culture and animal models. This phenomenon of photobiomodulation has been applied clinically in the treatment of soft tissue injuries and the acceleration of wound healing. The mechanism of photobiomodulation by red to near-IR light at the cellular level has been ascribed to the activation of mitochondrial respiratory chain components, resulting in initiation of a signaling cascade that promotes cellular proliferation and cytoprotection. A growing body of evidence suggests that cytochrome oxidase is a key photoacceptor of light in the far-red to near-IR spectral range.

Cytochrome oxidase is an integral membrane protein that contains four redox active metal centers and has a strong absorbance in the far-red to near-IR spectral range detectable in vivo by near-IR spectroscopy.

Moreover, 660–680 nm of irradiation has been shown to increase electron transfer in purified cytochrome oxidase, increase mitochondrial respiration and ATP synthesis in isolated mitochondria, and up-regulate cytochrome oxidase activity in cultured neuronal cells.

LED photostimulation induces a cascade of signaling events initiated by the initial

absorption of light by cytochrome oxidase. These signaling events may include the activation of immediate early genes, transcription factors, cytochrome oxidase subunit gene expression, and a host of other enzymes and pathways related to increased oxidative metabolism.

In addition to increased oxidative metabolism, red to near-IR light stimulation of mitochondrial electron transfer is known to increase the generation of reactive oxygen species. These mitochondrially generated reactive oxygen species may function as signaling molecules to provide communication between mitochondria and the cytosol and nucleus129.

Fujimaki et al in 2003 conducted a study to examine the effects of low-level laser therapy (LLLT) on production of reactive oxygen (ROS) species by human neutrophils and found that the attenuating effect of LLLT was larger in neutrophils of smokers than non-smokers, while the amount of produced ROS was larger in neutrophils of smokers. Expression of CD11b and CD16 on neutrophil surface was not affected by LLLT. This proves that there is a possible usage of LLLT to improve wound healing in smokers130.

Low-level laser therapy (LLLT) has been promoted for its beneficial effects on tissue healing and pain relief. However, according to the results of in vivo studies, the

effectiveness of this modality varies. A study was done by  Hopkins et al in 2004 to assess the putative effects of LLLT on healing using an experimental wound model.

Two standardized 1.27-cm2 abrasions were induced on the anterior forearm. After wound cleaning, standardized digital photos were recorded. Each subject then received LLLT (8 J/cm2; treatment time 5 2 minutes, 5 seconds;pulse rate 5,700 Hz) to 1 of the 2 randomly chosen wounds from either a laser or a sham 46-diode cluster head. At days 6, 8, and 10, follow-up testing revealed that the laser group had smaller wounds than the sham group for both the treated and the untreated wounds.

The LLLT resulted in enhanced healing, as measured by wound contraction.  Data indicates that LLLT is an effective modality to facilitate wound contraction of partial-thickness wounds131.

José Cláudio et al in 2006 selected twenty patients with periodontal disease, in whom treatment was planned for gingivectomy to bilateral maxillary and mandibular premolar teeth. After surgery, one side was submitted to LLLT using a 685-nm wavelength, output power of 50 mW, and energy density of 4 J/cm2. The other side was used as the control and did not receive laser irradiation. Healing was evaluated, clinically and biometrically, immediately post-surgery and at days 3, 7, 14, 21, 28, and 35. Clinical evaluation showed better repair for the laser group, mainly after the third

day. LLLT proved as an effective adjunctive treatment that appeared to promote healing following gingivectomy132.

Juliano et al in 2008 conducted a study to histometrically evaluate the influence of photodynamic therapy on bone loss in furcation areas in rats with experimentally induced periodontal disease.Ligatures were placed on the first mandibular molar in rats. Animals were divided into four groups: control group = no treatment; methylene blue group (MB) = treated topically with methylene blue (100 μg/ml); laser group (LLLT) = treated with low-level laser therapy; and photodynamic therapy group (PDT) = treated topically with MB followed by LLLT (4.5 J/cm2).

Rats from all groups were sacrificed at 7, 15, or 30 days postoperatively. The area of bone loss in the furcation region of the first molar was histometrically analyzed. The PDT group demonstrated less bone loss compared to the other groups at 7 days (1.986 ± 0.417 mm2); at 15 days, the PDT (1.641 ± 0.115 mm2) and MB groups (1.991 ± 0.294 mm2) demonstrated less bone loss compared to the control (4.062 ± 0.416 mm2) and LLLT (2.641 ± 0.849 mm2) groups.

Within the parameters used in this study, PDT may be an effective alternative for control of bone loss in furcation areas in periodontitis133.

DESTRUCTION OF BACTERIA BY LETHAL LASER

PHOTOSENSITIZATION

High-power heat-generating lasers, such as carbon dioxide and Nd:YAG surgical lasers, have well recognized destructive effects on bacteria, and this has led to the development of techniques for sterilizing wounds, cavity preparations and root canals. With heat-generating lasers, high irradiances (typically 100 J) are required for most ‘sterilizing’ applications, and data from in vivo studies indicate that risks of thermal injury can be substantial, although this varies markedly according to the wavelength used. In contrast, if low power laser energy could be coupled (using dyes) into the bacterial cell wall, the energy required for destruction of bacteria would be quite small, and would pose little if any risk of injury to the pulp and periodontal ligament.

The term lethal laser photosensitization (LLP) refers to the process whereby laser radiation emitted from a low power laser device activates a dye which in turn exerts a lethal effect on particular cells, such as bacteria. LLP is a specific interaction, in that treatment with the laser alone (i.e., in the absence of the enhancing dye), or with the dye alone, produces minimal effect. Similarly, treatment with laser followed by dye is without benefit. However, when exposure to the dye is followed by laser irradiation , there are changes within the dye which lead to dramatic chemical events within the target, with loss of viability as the end result.

Since LLP relies on the use of the dye as an optically absorbing agent (or chromophore), it is essential to match the laser wavelength used with the chromophore, as well as to ensure that the chromophore itself does not exert toxic or irritant effects on the tissues. LLP is similar in principle to photodynamic therapy, in which laser energy activates otherwise non-toxic dyes producing reactive oxygen species that cause injury and death of tumour cells.

Several laboratory studies have explored applications of LLP in dentistry. The visible red light emitted from a HeNe gas laser (wavelength 632 nm) can exert an inhibitory action on dental plaque micro-organisms when an appropriate dye is used.

Streptococcus sobrinus is sensitive to LLP under these conditions, while Escherichia coli, a Gram-negative enteric micro-organism, is resistant. The effect on the cariogenic micro-organism S.sobrinus is rapid, with leakage of the bacterial membrane detectable after 2 minutes of laser exposure. The bactericidal action of LLP can be obtained with a range of blue, purple, and green dyes within the phenylmethane family, all of which are strong absorbers at 632 nm. Similar findings were obtained in a later study, which demonstrated destruction of biofilms of Streptococcus sanguis, Porphyromonas gingivalis, Fusobacterium nucleatum and Actinobacillus actino - mycetemcomitans on agar plates by LLP. Using a 7.3 mW HeNe gas laser, either toluidine blue O or methylene blue resulted in killing of all four target organisms after lasting for 30 seconds.

These blue dyes are known to serve as potent sensitizers for a range of bacteria, for laser emissions in the visible red portion of the spectrum. In a more recent study, killing via LLP has been achieved for the cariogenic bacteria Streptococcus mutans, S. sobrinus, Lactobacillus casei and Actinomyces viscosus, using toluidine blue.

For each organism, suspensions containing 106 colony forming units could be inactivated by irradiation for 60 seconds. It has since been demonstrated that longer exposure times can result in kills exceeding 107 colony forming units. These investigations, conducted using cultures on agar plates, provide a clue as to the possible value of LLP for destruction of cariogenic bacteria. However, there are several important limitations of these studies which should be recognized. First, the experimental situations used have not replicated those encountered clinically, and the problems of bacteria present on smear layers or in dentinal tubules have not been addressed. As well, the use of HeNe gas lasers as a source of laser energy is less than ideal for a clinical application because these devices require optical systems and power supplies which are complex. Their conversion efficacy (electrical energy to laser energy) is not high.

By comparison, indium:gallium:arsenide:phosphorus (InGaAsP) semiconductor diode lasers generate emissions of similar wavelength to HeNe lasers; however, their smaller size, higher efficiency, low cost, and more versatile nature makes them

better suited for LLP. Low power versions of these devices are used in lecture theatre pointers.

There is considerable current interest in developing LLP techniques using visible diode lasers. Work in the author’s laboratory is evaluating the feasibility of using InGaAsP laser diodes operating in the visible spectrum (wavelength 670 nm) in conjunction with methylene blue to sterilize a carious cavity and infected dentinal tubules containing a known ‘bio-load’ of S. mutans. This would have obvious clinical relevance in terms of managing deep dentinal carious lesions. With the ability to rapidly and effectively sterilize the floor of deep carious lesions without significant heat, more conservative approaches to the removal of demineralized and infected tooth structure could be used.

Of note, LLP is able to exert significant effects even when the cariogenic organisms are protected in a matrix of demineralized dentine. As mentioned previously, laser radiation in the red visible and near infrared region is able to pass through demineralised dentine. While LLP techniques can still be applied effectively to reach micro-organisms enmeshed in dentine, higher irradiances are required to exert the same destructive effect than when the microorganisms are on the surface (4.8 vs. 1.2 J). To achieve these irradiances in a clinically acceptable time frame requires the use of higher power diode lasers and more efficient delivery systems, and current work is addressing these issues.

Lethal laser photosensitization technology may also have application in the treatment of hypersensitive cervical dentine covered by dental plaque, where the technique would be used in tandem with desensitization of the dental pulp by LLLT. In terms of managing plaque on hypersensitive dentine, LLP has recently been shown to be a viable technique for killing bacteria in supragingival plaque. Following LLP with either HeNe or GaAlAs lasers, substantial reductions were demonstrated in the total anaerobic count as well as in the number of viable streptococci and actinomyces present in the samples. An irradiance of only 1.3 J resulted in log10 reductions of 3.0, 5.4 and 3.3 in the total anaerobic count, streptococci and actinomyces respectively.

In addition to its use in the management of dental caries, the LLP technique has potential application for the destruction of microorganisms in the root canal system and in periodontal pockets, as an adjunct to conventional techniques. However, much further work is required to develop protocols which could be used in a clinical setting and to improve the laser delivery systems to permit access to all areas of the mouth128.

LASER REGULATIONS AND SAFETY

SAFETY WITH INDUSTRIAL LASERS (BEAM & NON-BEAM HAZARDS)

Potential hazards related to the use of lasers can generally be divided into primary and secondary hazards. The laser beam itself represents the primary potential hazard, as it can affect humans or objects – in the form of raw beam, focussed beam, directly reflected beam, or scattered radiation.

Secondary potential hazards are further subdivided into direct and indirect hazards:

- direct potential hazards are caused by technical components of the laser installation (high voltage, excitation radiation, laser gases, optics),

- indirect potential hazards are generated by the interaction of the laser beam with materials or the atmosphere.

Indirect potential hazards include the UV-radiation caused by plasma formation, hazardous substances generated during material processing, and also potential ignition of explosive materials and the danger of fire.

Laser Radiation

There are two potential dangers to personnel associated with laser radiation: eye and skin damage. The factors that can contribute to tissue injury and influence the degree of damage from laser beam exposures include:

a) Wavelength of laser radiation.

b) Tissue spectral absorption, reflection and transmission.

c) Strength of irradiance of incident laser beam.

d) Size of irradiated area.

e) Exposure duration.

f)  Pupil size.

g) Location of retinal injury.

h) Laser pulse characteristics.

Depending on wavelength, damage to either the cornea, retina, or both, of the eye is possible. Exposure to radiation from a CO2 laser (10.6 µm) typically results in corneal damage. The radiation from an Nd:YAG laser, at 1.06 µm, is much closer to the visible spectrum (400 to 700 nm) and can be transmitted by the cornea and lens. The lens will focus the laser light on the retina, causing severe and permanent damage to the retina and other intraocular material. This focusing, by the lens of the eye, can cause even low-power diffuse laser light to be focussed to a sufficient power density to cause retinal damage. Low-power helium-neon (He-Ne) lasers (0.633 µm), often used for alignment purposes, may also present a hazard.

Skin damage is restricted primarily to burns. It should be noted, however, for high-powered lasers these burns can be deep and cause severe and permanent damage.

The most effective prevention of injury is to ensure that the laser beam is encapsulated so that no human exposure can occur. For the case of an exposure to the beam, the level of exposure will determine if injury occurs. The level of exposure or irradiance which can be thought of as the border between safe and potentially harmful is called “Maximum Permissible Exposure”, MPE.

The area around a laser installation where the MPEs for the eye can be exceeded is called the “Nominal Ocular Hazard Area”, NOHA. The corresponding distance from the laser exit aperture is called the “Nominal Ocular Hazard Distance”, NOHD. The NOHD depends on the laser power, geometrical laser beam parameters such as the divergence, and the MPE. MPE values for the eye and the skin as a function of wavelength and exposure duration are listed in AS/NZS 2211.1. The standard also contains a number of sample NOHD calculations for a range of geometries and exposure situations.

As MPE and NOHD evaluations are quite complicated and involved, the laser classification scheme outlined earlier is used for the hazard evaluation process.

For instance for Class 1 lasers, the exposure will always be below the MPE.

Class 3 B lasers emit radiation which is significantly above the MPE for eye, and for Class 4 also above the MPE for the skin, however, it depends on the beam geometry,

set-up and application, if this hazard exists only close to the exit aperture or if it extends over several kilometres.

Eye Protection

Information on eye protectors suitable for use with particular lasers and operations together with their required marking is given in BS EN207 Personal eye-protection. Filters and eye-protectors against laser radiation (laser eyeprotectors) and BS EN208 Personal eye protection. Eye-protectors for adjustment work on lasers and laser systems (laser adjustment eye-protectors).

The following should be considered when specifying suitable protective eyewear:

a) Wavelength(s) of operation.

b) Radiant exposure or irradiance.

c) Maximum permissible exposure (MPE).

d) Optical density of eyewear at laser output wavelength.

e) Visible light transmission.

f)  Radiant exposure or irradiance at which damage to eyewear occurs.

g) Need for prescription glasses.

h) Comfort and ventilation.

i)  Degradation or modification of absorbing media, even if temporary or transient.

j)  Strength of materials (resistance to shock).

k) Peripheral vision requirements.

l)  Any relevant legislation.

A method of selecting the laser or laser adjustment eye protector most suited to the hazards associated with the use of a particular laser is given in AS/NZS 1336

Recommended practices for occupational eye protection :

Special attention has to be given to the resistance and stability against laser radiation when choosing eyewear for Class 4 laser products.

Laser eye protectors and laser adjustment eye protectors which have been damaged or have undergone a colour change must not be used. Eye protectors are only intended to protect against accidental exposure as the ratings are based on a maximum exposure of 10s (for a continuous wave laser) or 100 pulses (for a pulsed laser). Eye protectors are not intended to be used for looking directly into the beam.

A precaution must be added here – standard safety glasses alone do not provide protection. Any laser eyewear, plain or prescription, must be labelled in accordance with BS EN207 or BS EN208 with information adequate to ensure the proper choice of eyewear with particular lasers. In some laser systems, ultraviolet light may be leaked into the workplace. Thus the eyewear should provide primary beam protection, secondary radiation protection and also ultraviolet protection.

Skin protection

Where personnel may be exposed to levels of radiation that exceed the MPE for the skin, suitable clothing should be worn. Class 4 laser products present a potential fire hazard and protective clothing worn should be made from a suitable flame and heat-resisting material. Special attention must be given to resistance and long-term stability against laser radiation when choosing protective clothing for use with Class 4 laser products.

Electrical Hazards

The voltages used in lasers are sufficient to cause fatal injuries to personnel and account for most laser-related fatalities. All electrical equipment associated with laser beam materials processing should be installed in conformance to AS/NZS 3000 Electrical installations (known as the Australian / New Zealand Wiring Rules).

All doors and access panels should be properly secured, either electrically or mechanically, to prevent access by unauthorised personnel to electrical components, especially those operating at the laser excitation potential. All personnel working on or around high-voltage components should be trained in the proper safety techniques for electrical systems, as well as in the technique of removing a victim from an electrical circuit and administering cardiopulmonary resuscitation (CPR). Personnel

should be aware of and adhere to any additional electrical safety requirements of the laser system installed in their facility.

Usually, the best source of safety information is provided in the instruction manual from the manufacturer of the laser system. Always read, understand and follow the manufacturer’s recommended safety procedures.

Fumes and Gases

Welding, cutting and drilling, and surface modification with lasers may result in the generation of fumes, dust, and gases that can be hazardous to personnel.

These airborne contaminants may include:

a) Vapourised target material and reaction products in the form of metal particles and oxides.

b) Gases from the flowing gas laser systems or from the by-products of laser reactions, such as ozone, nitrous oxide, carbon monoxide and carbon dioxide.

c) Gases or vapours from cryogenic coolants.

d) Gases used to assist laser-target interactions, such as oxygen.

The hazards associated with welding and cutting of metals have been documented in a variety of publications. It should be noted that some organic materials, such as plastics, can generate fumes that are hazardous. Care should be taken to avoid the  excessive build-up of laser discharge gases, shielding gases, and assist gases, especially in enclosed spaces where oxygen can be displaced. All necessary environmental engineering measures for fume and gas control (external venting, filtering, etc) should be taken to prevent the accidental inhalation of harmful

concentrations of fumes and gases by personnel working on or around laser materials processing equipment.

Exhaust of these fumes may violate local or federal standards, and implications should be considered before using equipment. The possible toxicity of the workpiece and consumables (wire, powder, etc) should be determined before laser beam materials processing begins. Also, for all materials, the Material Safety Data Sheet (MSDS), available from the material supplier, should be consulted to determine what hazards exist.

Fire Hazard

Since the laser system produces a very small spot size with high energy, the hazard of fire is present if the beam hits flammable material. Keep flammables away from the welding or cutting area. Be sure to cover and protect anything flammable in the area since reflected radiation could start fires in unexpected areas.

The potential for explosions at the capacitor bank or optical pump systems exist during the operation of some high-power laser systems. Metallic and non-metallic dusts such as airborne particles coming from the target area in the laser cutting, drilling and welding operations may be capable of causing fire or explosion. Explosive reactions of chemical laser reagents or other chemicals or gases are also possible.

Secondary Radiation Hazards

Viewing of the visible radiation emitted during laser material processing can also be harmful to eyesight. During welding, a bright plume, similar in appearance to a welding arc, is generated from the interaction between the laser beam and material being processed. The size and intensity of this plume is a function of the material being processed, the power level, and the shielding gas used.

Consequently, no exact guidelines can be given. However, the radiation emitted is broad-band, from the ultraviolet, which can cause sunburn and arc eye through the visible to the infrared, which can contribute to the formation of cataracts.

As the plume is generally too bright for direct viewing, adequate filtering, such as welding shades, should be employed for eye protection. As a general guideline, the filter used should be of sufficient optical density to ensure the viewer’s comfort at the

highest level of light intensity encountered and there should be no evidence of eye irritation after exposure.

The optical viewing system should provide filtering in conformance to AS/NZS 1337 and 1338 and should include provisions for filtering the visible and ultraviolet radiation from the plume as well as the laser radiation. All persons involved with laser beam materials processing should be instructed in the use of proper optical filtering and should be required to use such protection.

Optical fibre system associated hazards

There are many Nd:YAG lasers on the market now where fibre beam delivery is the only method of energy delivery. The optical fibre alone should not be relied on as a safe enclosure. Hazardous situations associated with the optical fibre itself include:

- Leakage due to inappropriate acceptance angle.

- Leakage due to the fibre being subject to too tight a bending radius.

- Fibre breakage due to excessive bending or twisting. If the fibre is carrying

laser power at the time the break occurs then, because of the relatively small fibre diameter, the laser beam will emerge from what is effectively a high

intensity point source as a divergent beam.

- Leakage and excess heat generation due to the ingress of dirt or other contamination onto the fibre input or output faces.

CONTROL MEASURES FOR LASER INSTALLATIONS

Laser systems used for material processing can cause injury from both the direct beam

or its specular reflections and from diffuse reflections. They also present associated hazards as outlined above. The following control measures should be employed to minimise these risks:

a) The laser should only be operated in a controlled area. Class 4 lasers should be operated by remote control, whenever practicable, thus eliminating the need for personnel to be physically present in the laser environment.

b) The entrances to areas should be posted with a standard laser warning sign as specified in AS/NZS 2211.1.

c) Beam paths should be enclosed whenever practicable. Access to the laser environment during laser operation should be limited to persons wearing proper laser protective eyewear and protective clothing. Beam paths should avoid work area where possible, and long sections of tubes should be mounted so that thermal expansion, vibration, and other sources of movements in them do not significantly affect the alignment of beam forming components.

d) Good room illumination is important in areas where laser eye protection is worn. Light-coloured diffuse wall surfaces help to achieve this condition.

e) Fire, thermally induced aberrations in optical components and the melting or vaporisation of solid targets designed to contain the laser beam, are all potential hazards induced by the radiation of Class 4 lasers. A suitable beam stop should be provided. Materials which may be degraded by the beam must be avoided.

f) Special precautions may be required to prevent unwanted reflections, especially for infrared laser radiation, and the beam and target area should be surrounded by a material opaque to the laser wavelength. (Even dull metal surfaces may become highly specular at the CO2 wavelength of 10.6 µm). Local screening should be used wherever practicable to reduce the extent of reflected radiation.

g) The alignment of optical components in the path of a Class 4 laser beam should

be initially and periodically checked.

h) The remote interlock connector should be connected to an emergency master

disconnect interlock or to room, door or fixture interlocks. The person in charge

may be permitted momentary override of the remote interlock connector to allow

access to other authorised persons if it is clearly evident that there is no optical

radiation hazard at the time and point of entry.

i) Laser systems not in use should be protected against unauthorised use by removal of the key of the key control.

Medical Surveillance

Laser workers whose work involves a significant risk of exposure to laser radiation in

excess of the Maximum Permissible Exposure (MPE) should have eye examinations

and, where appropriate, skin examinations carried out at commencement and termination of the position. For anyone at increased risk of laser damage, more frequent eye examinations may be prudent.

Laser Safety Officer

The “Laser Safety Officer” is defined in AS/NZS 2211.1 as “one who is competent in the evaluation and control of laser hazards and has responsibility for oversight of the control of laser hazards”.

For each organisation using potentially harmful laser systems, management should assign an individual to serve as LSO to be responsible for their safe use. The LSO

requires this authority from management to ensure users maintain safe practices with their equipment. Management may employ an LSO to oversee the responsibility while relying on technical input from an outside expert. The LSO generally receives special safety training and has access to laser safety guidance documents, equipment and support staff commensurate with the extent of his or her responsibilities.

The LSO undertakes a risk assessment before a laser or laser system is first used, or after it has been modified. The LSO determines potentially hazardous areas, degree of hazard, necessary administrative and engineering controls, and necessary safety instructions, and determines appropriate personnel protective equipment such as optical density (OD) for laser protective eyewear. The LSO also maintains an inventory of lasers with detailed information such as: manufacturer, model number, quantity, physical location, user organisation, laser active medium, laser radiant power or energy, laser radiant wavelength(s), laser application, and laser hazard category.

Other useful information (eg history and last maintenance date) may be required by an individual LSO134.

CONCLUSION

A reliable procedure for laser application in nonsurgical periodontal therapy should be established by further studies, and clinicians should follow the results of scientific investigations to obtain successful outcomes.  Comprehensive treatment including preparation of both hard and soft tissue walls within pockets should be considered for more effective nonsurgical periodontal therapy in the future, and this is what may be accomplished by lasers.

As understanding of the nature of laser light evolves, lasers will be used more effectively in the treatment of periodontal diseases. Currently, among the different types of lasers available, Er:YAG and Er,Cr:YSGG are the most promising ones which possess characteristics suitable for dental treatment, due to their dual ability to ablate soft and hard tissues with minimal damage.

Laser systems applying the ablation effect of light energy, which is completely different from conventional mechanical debridement, may emerge as a new technical modality for nonsurgical periodontal therapy and will have the potential to become a part of routine dental care in the near future.

BIBLIOGRAPHY

- Lasers in Dentistry By: Dr. David Hornbrook : 2006.

- The first laser Charles H. Townes from A Century of NaturE :Twenty-One Discoveries that  Changed Science and the World:2005.

- Lasers in dentistry;Leo J. Miserendino, DDS, MS; Robert M Pick, DDS, MS, FACD, FICD;1995, Quintessence Publishing.

- At http://www.kigre.com/files/historylasers.pdf

- At http://www.bell-labs.com/history/laser

- Introduction, history of lasers and laser light production. British dental journal 2007;202: 21-31.

- Dental Clinics of North America 48 (2004) ; 773,774

- Laser safety – April 2003;http://www.wtia.com.au

- http://www.geocities.com/muldoon432/Laser Types and Classifications. htm

- Laser-tissue interaction. British dental journal 2007;202:73-81.

- Lasers in non surgical periodontal therapy. Periodontology 2000;36,2004,59-97.

- Application of Laser in Periodontics : A New Approach in Periodontal Treatment ; Dr. Dae-hyun Lee October 2007 vol.12 no.10.

- Arashiro DS, Rapley JW, Cobb CM, Killoy WJ. Histologic evaluation of porcine skin incisions produced by CO2 laser, electrosurgery, and scalpel. Int J Periodontics Restorative Dent 1996:16:479–491.

- Rossmann JA, Gottlieb S, Koudelka BM, McQuade MJ. Effects of CO2 laser irradiation on gingiva. J Periodontol 1987:58:423–425.

- Barak S, Kaplan I. The CO2 laser in the excision of gingival hyperplasia caused by nifedipine. J Clin Periodontol 1988:15:633–635.

- Pick RM, Pecaro BC. Use of the CO2 laser in soft tissue dental surgery. Lasers Surg Med 1987:7:207–213.

- Pick RM, Pecaro BC, Silberman CJ. The laser gingivectomy. The use of the CO2 laser for the removal of phenytoin hyperplasia. J Periodontol 1985: 56: 492–496.

- Israel M, Cobb CM, Rossmann JA, Spencer P. The effects of CO2, Nd:YAG and Er:YAG lasers with and without surface coolant on tooth root surfaces. An in vitro study. J Clin Periodontol 1997:24:595–602.

- Sasaki KM, Aoki A, Ichinose S, Ishikawa I. Morphological analysis of cementum and root dentin after Er:YAG laser irradiation. Lasers Surg Med 2002:31:79–85.

- Sasaki KM, Aoki A, Masuno H, Ichinose S, Yamada S, Ishikawa I. Compositional analysis of root cementum and dentin after Er:YAG laser irradiation compared with CO2 lased and intact roots using Fourier transformed infrared spectroscopy. J Periodontal Res 2002:37:50–59.

- Shariati S, Pogrel MA, Marshall GW Jr, White JM, Powell GL, Morton TH, Whisenant BK. Structural changes in dentin induced by high energy, continuous wave carbon dioxide laser. Lasers Surg Med 1993:13:543–547.

- Spencer P, Cobb CM, McCollum MH, Wieliczka DM. The effects of CO2 laser and Nd:YAG with and without water/air surface cooling on tooth root structure: correlation between FTIR spectroscopy and histology. J Periodontal Res 1996:31:453–462.

- Tucker D, Cobb CM, Rapley JW, Killoy WJ. Morphologic changes following in vitro CO2 laser treatment of calculus-ladened root surfaces. Lasers Surg Med 1996:18:150–156.

- Misra V, Mehrotra KK, Dixit J, Maitra SC. Effect of a carbon dioxide laser on periodontally involved root surfaces. J Periodontol 1999:70:1046–1052.

- Barone A, Covani U, Crepsi R, Romanos GE. Root surface morphological changes after focused versus defocused CO2 laser Irradiation: a scanning electron microscopy analysis. J Periodontol 2002:73:370–373.

- Crespi R, Barone A, Covani U, Ciaglia RN, Romanos GE. Effects of CO2 laser treatment on fibroblast attachment to root surfaces. A scanning electron microscopy analysis.J Periodontol 2002:73:1308–1312.

- Miyazaki A, Yamaguchi T, Nishikata J, Okuda K, Suda S, Orima K, Kobayashi T, Yamazaki K, Yoshikawa E, Yoshie H. Effects of Nd:YAG and CO2 laser treatment and ultrasonic scaling on periodontal pockets of chronic periodontitis patients. J Periodontol 2003:74:175–180.

- Midda M, Renton-Harper P. Lasers in dentistry. Br Dent J 1991: 170: 343–346.

- Myers TD, Murphy DG, White JM, Gold SI. Conservative soft tissue management with the low-powered pulsed Nd:YAG dental laser. Pract Periodontics Aesthet Dent 1992:4:6–12.

- Pick RM, Colvard MD. Current status of lasers in soft tissue dental surgery. J Periodontol 1993:64:589–602.

- Romanos GE. Clinical applications of the Nd:YAG laser in oral soft tissue surgery and periodontology. J Clin Laser Med Surg 1994:12:103–108.

- White JM, Fagan MC, Goodis HE. Intrapulpal temperatures during pulsed Nd:YAG laser treatment of dentin in vitro. J Periodontol 1994, 65: 255–259.

- Sulewski JG. Historical survey of laser dentistry. Dent Clin North Am 2000: 44:717–752.

- Bassi G, Chawla S, Patel M. The Nd:YAG laser in caries removal. Br Dent J 1994:177:248–250.

- White JM, Goodis HE, Setcos JC, Eakle S, Hulscher BE, Rose CL. Effects of pulsed Nd:YAG laser energy on human teeth: a three-year follow-up study. J Am Dent Assoc 1993:124: 45–51.

- Sulewski JG. Historical survey of laser dentistry. Dent Clin North Am 2000: 44:717–752.

- Tseng P, Liew V. The potential applications of a Nd:YAG dental laser in periodontal treatment. Periodontology (Australia) 1990:11:20–22.

- Tseng P, Liew V. The use of a Nd:YAG dental laser in periodontal therapy. Aust Dent Assoc News Bull 1991: Nov:3–6.

- Tucker D, Cobb CM, Rapley JW, Killoy WJ. Morphologic changes following in vitro CO2 laser treatment of calculus - ladened root surfaces. Lasers Surg Med 1996:18:150–156.

- Arcoria CJ, Vitasek-Arcoria BA. The effects of low-level energy density Nd:YAG irradiation on calculus removal. J Clin Laser Med Surg 1992:10: 343–347.

- Radvar M, Creanor SL, Gilmour WH, Payne AP, McGadey J, Foye RH, Whitters CJ, Kinane DF. An evaluation of the effects of an Nd:YAG laser on subgingival calculus, dentine and cementum. An in vitro study. J Clin Periodontol 1995:22:71–77.

- Israel M, Cobb CM, Rossmann JA, Spencer P. The effects of CO2, Nd:YAG and Er:YAG lasers with and without surface coolant on tooth root surfaces. An in vitro study. J Clin Periodontol 1997:24:595–602.

- Morlock BJ, Pippin DJ, Cobb CM, Killoy WJ, Rapley JW. The effect of Nd:YAG laser exposure on root surfaces when used as an adjunct to root planing: an in vitro study. J Periodontol 1992:63:637–641.

- Spencer P, Cobb CM, McCollum MH, Wieliczka DM. The effects of CO2 laser and Nd:YAG with and without water/ air surface cooling on tooth root structure: correlation between FTIR spectroscopy and histology. J Periodontal Res 1996:31:453–462.

- Spencer P, Trylovich DJ, Cobb CM. Chemical characterization of lased root surfaces using Fourier transform infrared photoacoustic spectroscopy. J Periodontol 1992:63:633–636

- Dowker SE, Elliott JC. Infrared absorption bands from NCO– and NCN2– in heated carbonate-containing apatites prepared in the presence of NH4 + ions. Calcif Tissue Int 1979:29:177–178.

- Dowker SE, Elliott JC. Infrared study of the formation, loss, and location of cyanate and cyanamide in thermally treated apatites. J Solid State Chem 1983: 49:334–340.

- Thomas D, Rapley J, Cobb C, Spencer P, Killoy W. Effects of the Nd:YAG laser and combined treatments on in vitro fibroblast attachment to root surfaces. J Clin Periodontol 1994:21:38–44.

- Wilder-Smith P, Arrastia AM, Schell MJ, Liaw LH, Grill G, Berns MW. Effect of ND:YAG laser irradiation and root planing on the root surface: structural and thermal effects. J Periodontol 1995:66:1032–1039.

- Fukuda M, Minoura S, Ishikawa K, Ogura N, Ueda N, Murase M, Sugihara N, Kato K, Nakagaki H, Noguchi T. Effects of Nd:YAG laser irradiation on endotoxin in exposed cementum [in Japanese, English abstract]. Jpn J Conserv Dent 1994:37:711–716.

- Liu CM, Shyu YC, Pei SC, Lan WH, Hou LT. In vitro effect of laser irradiation on cementum-bound endotoxin isolated from periodontally diseased roots. J Periodontol 2002:73:1260–1266

- Cobb CM, McCawley TK, Killoy WJ. A preliminary study on the effects of the Nd:YAG laser on root surfaces and subgingival microflora in vivo. J Periodontol 1992:63:701– 707.

- Gold SI, Vilardi MA. Pulsed laser beam effects on gingiva. J Clin Periodontol 1994:21:391–396.

- Neil ME, Mellonig JT. Clinical efficacy of the Nd:YAG laser for combination periodontitis therapy. Pract Periodontics Aesthet Dent 1997:9:1–5.

- Miyazaki A, Yamaguchi T, Nishikata J, Okuda K, Suda S, Orima K, Kobayashi T, Yamazaki K, Yoshikawa E, Yoshie H. Effects of Nd:YAG and CO2 laser treatment and ultrasonic scaling on periodontal pockets of chronic periodontitis patients. J Periodontol 2003:74:175–180.

- White JM, Gekelman D, Shin K, Park J, Swenson T, Rouse B, Tran K, Bullard S, Scott-Beckles B, Oto MJSB, Yamamoto A. Laser interaction with dental soft tissues:What do we know from our years of applied scientific research? Proc SPIE 2002:4610:39–48.

- Coluzzi DJ. Lasers and soft tissue curettage: An update. Compendium 2002: 23:1104–1111.

- Gutknecht N, Zimmermann R, Lampert F. Lasers in periodonology: state of the art. J Oral Laser Appl 2001:1:169–179.

- Hale GM, Querry MR. Optical constants of water in the 200-nm to 200-µm wavelength region. Appl Opt 1973:12:555–563.

- Robertson CW, Williams D. Lambert absorption coefficients of water in the infrared. J Opt Soc Am 1971:61:1316–1320

- Featherstone JDB. Caries detection and prevention with laser energy. Dent Clin North Am 2000:44:955–969.

- Cozean C, Arcoria CJ, Pelagalli J, Powell GL. Dentistry for the 21st century? Erbium:YAG laser for teeth. J Am Dent Assoc 1997:128:1080–1087.

- Keller U, Hibst R. Kariestherapie mit dem Erbium:YAG Laser – Erste Klinische Ergebnisse [in German; English abstract]. ZWR 1993:102:804–807.

- Keller U, Hibst R. Effects of Er:YAG laser in caries treatment: a clinical pilot study. Lasers Surg Med 1997:20: 32–38.

- Kumazaki M. Removal of hard dental tissue (cavity preparation) with the Er:YAG laser. In: Loh HS, editor. 4thInternational Congress on Lasers in Dentistry. Bologna:Monduzzi Editore, 1995:151–157.

- Matsumoto K, Nakamura Y, Mazeki K, Kimura Y. Clinical dental application of Er:YAG laser for Class V cavity preparation. J Clin Laser Med Surg 1996: 14:123–127.

- Aoki A, Watanabe H, Namiki N, Takiguchi T, Miyazawa Y, Suzuki M, Hasegawa K, Ishikawa I. Periodontal soft tissue management with a high pulse rate Er:YAG laser. In: Ishikawa I, Frame JW, Aoki A, editors. Lasers in Dentistry – Revolution of Dental Treatment in the New Millennium. Amsterdam: Elsevier Science, 2003:367–369.

- Watanabe H, Ishikawa I, Suzuki M, Hasegawa K. Clinical assessments of the erbium:YAG laser for soft tissue surgery and scaling. J Clin Laser Med Surg 1996:14: 67–75.

- Sulewski JG. Historical survey of laser dentistry. Dent Clin North Am 2000: 44:717–752.

- Walsh JT Jr, Flotte TJ, Deutsch TF. Er:YAG laser ablation of tissue: effect of pulse duration and tissue type on thermal damage. Lasers Surg Med 1989:9: 314–326.

- Burkes EJ Jr, Hoke J, Gomes E, Wolbarsht M. Wet versus dry enamel ablation by Er:YAG laser. J Prosthet Dent 1992:67:847–851.

- Paghdiwala AF. Er:YAG laser hard tissue effects. In: Moretti M, editor. Laser in Dentistry. Massachusetts:PennWell Publishing, 1991:63–75.

- Visuri SR, Walsh JT, Wigdor HA. Erbium laser ablation of dental hard tissue: Effect of water cooling. Lasers Surg Med 1996:18:294–300.

- Featherstone JDB. Caries detection and prevention with laser energy. Dent Clin North Am 2000:44:955–969.

- Hibst R, Keller U. Experimental studies of the application of the Er:YAG laser on dental hard substances. I. Measurement of the ablation rate. Lasers Surg Med 1989:9:338–344.

- Keller U, Hibst R. Experimental studies of the application of the Er:YAG laser on dental hard substances. II. Light microscopic and SEM investigations. Lasers Surg Med 1989:9:345–351.

- Koort HJ, Frentzen M. Laser effects on dental hard tissue. In: Miserendino, LJ, Pick, RM, editors. Lasers in Dentistry.Chicago: Quintessence, 1995:57–70.

- Fried D. IR laser ablation of dental enamel. Proc SPIE 2000:3910:136–148.

- Aoki A, Ando Y, Watanabe H, Ishikawa I. In vitro studies on laser scaling of subgingival calculus with an erbium: YAG laser. J Periodontol 1994:65:1097–1106.

- Schwarz F, Sculean A, Berakdar M, Georg T, Becker J. In vivo and in vitro effects of an Er:YAG laser, a GaAlAs diode laser and scaling and root planing on periodontally diseased root surfaces. A comparative histologic study. Lasers Surg Med 2003:32:359–366.

- Frentzen M, Braun A, Aniol D. Er:YAG laser scaling of diseased root surfaces. J Periodontol 2002:73:524–530.

- Ishikawa I. Letter to the editor. Re: Er:YAG laser Scaling of diseased root surfaces: Frentzen M, Braun A, Aniol D (J Periodontol 2002:73:524–530). J Periodontol 2002:73:1226.

- Aoki A, Miura M, Akiyama F, Nakagawa N, Tanaka J, Oda S, Watanabe H, Ishikawa I. In vitro evaluation of Er:YAG laser scaling of subgingival calculus in comparison with ultrasonic scaling. J Periodontal Res 2000:35:266–277.

- Israel M, Cobb CM, Rossmann JA, Spencer P. The effects of CO2, Nd:YAG and Er:YAG lasers with and without surface coolant on tooth root surfaces. An in vitro study. J Clin Periodontol 1997:24:595–602.

- Fujii T, Baehni PC, Kawai O, Kawakami T, Matsuda K, Kowashi Y. Scanning electron microscopic study of the effects of Er:YAG laser on root cementum. J Periodontol 1998:69:1283–1290

- Hibst R, Keller U, Steiner R. Die Wirkung gepulster Er:YAG Laserstrahlung auf Zahngewebe. Laser Med Surg 1988;4:163–165 (in German).

- Paghdiwala AF. Application of the erbium:YAG laser on hard dental tissues: Measurement of the temperature changes and depth of cut. In: Laser Institute of America Proceedings (International Congress on Applications of Lasers and Electro-Optics) 1989;64:192–201.

- Kayano T, Ochiai S, Kiyono K, Yamamoto H, Nakajima S, Mochizuki T. Effects of Er:YAG laser irradiation on human extracted teeth. Kokubyo Gakkai Zasshi 1989;56:381–392 (in Japanese).

- Kayano T, Ochiai S, Kiyono K, Yamamoto H, Nakajima S, Mochizuki T. Effect of Er:YAG laser irradiation on human extracted teeth. J Clin Laser Med Surg 1991;9:147–150.

- Kumazaki M. Removal of hard dental tissue (cavity preparation) with the Er:YAG laser. In: Hong-Sai L, ed. Proceedings of the 4th International Congress on Lasers in Dentistry. Singapore:International Congress on Lasers in Dentistry 1995:151–157.

- Ishikawa I, Aoki A, Watanabe H et al. Erbium:YAG laser, promising procedure for caries treatment. Dent Jpn 1997;33:165–169.

- Pelagalli J, Gimbel CB, Hansen RT, Swett A, Winn DW, 2nd. Investigational study of the use of Er:YAG laser versus dental drill for caries removal and cavity preparation- phase I. J Clin Laser Med Surg 1997;15:109–115.

- Ishikawa I, Sasaki KM, Aoki A, Watanabe H. Effects of Er:YAG laser on periodontal therapy. J Int Acad Periodontol 2003;5:23–28.

- Azzeh et al. Treatment of gingival hyperpigmentation by Erbium doped:Aluminium, Yttrium,Garnet laser for aesthetic purposes. Journal of Periodontol, January 2007

- Ishii S, Aoki A, Kawashima Y, Watanabe H, Ishikawa I. Application of an Er:YAG laser to remove gingival melanin hyperpigmentation– treatment procedure and clinical evaluation. J Jpn Soc Laser Dent 2002;13:89–96 (in Japanese).

- Kawashima Y, Aoki A, Ishii S, Watanabe H, Ishikawa I. Er:YAG laser treatment of gingival melanin pigmentation. Proceedings of the 8th International Congress on Lasers in Dentistry. 2003:245–248.

- Sato A, Ukon S, Hamano M, Miyoshi S. Gingival pigmentation due to prolonged retention of metal fragments. Jpn J Oral Biol 1997;39:565–571

- Aoki A, Ishii S, Mizutani K, Takasaki AA, Watanabe H, Ishikawa I. Application of the Er:YAG laser for removal of metal tattoo pigmentation originated from prosthetic metallic particles. The 15th Annual Meeting of the Japanese Society for Laser Dentistry. 2003:45 (Abstract) (In Japanese).

- Nelson JS, Orenstein A, LiawLH, Berns MW. Mid-infrared erbium. YAG laser ablation of bone: the effect of laser osteotomy on bone healing. Lasers Surg Med 1989;9:362–374.

100.Sasaki KM, Aoki A, Ichinose S, Ishikawa I. Ultrastructural analysis of bone tissue irradiated by Er:YAG laser. Lasers Surg Med 2002;31:322–332.

101.Williams TM, Cobb CM, Rapley JW,  Killoy WJ. Histologic evaluation of   alveolar bone following CO2 laser removal of connective tissue from periodontal defects. Int J Periodont Restor Dent 1995;15:497–506.

102.Sasaki KM, Aoki A, Ichinose S, Yoshino T, Yamada S, Ishikawa I. Scanning electron microscopy and Fourier transformed infrared spectroscopy analysis of bone removal using Er:YAG and CO2 lasers. J Periodontol 2002;73:643–652.

103.Mason ML. Using the laser for implant maintenance. Dent Today 1992;11:74 –75.

104.Matarasso S, Quaremba G, Coraggio F, Vaia E, Cafiero C, Lang NP. Maintenance of implants: an in vitro study of titanium implant surface modifications subsequent to the application of different prophylaxis procedures. Clin Oral Implants Res 1996;7: 64–72.

105.Kreisler M, Kohnen W, Marinello C et al. Bactericidal effect of the Er:YAG laser on dental implant surfaces: an in vitro study. J Periodontol 2002;73:1292–1298.

106.Matsuyama T, Aoki A, Oda S, Yoneyama T, Ishikawa I. Effects of the Er:YAG laser irradiation on titanium implant materials and contaminated implant abutment surfaces. J Clin Laser Med Surg 2003;21:7–17.

107.Schwarz F, Rothamel D, Becker J. Influence of an Er:YAG laser on the surface structure of titanium implants. Schweiz Monatsschr Zahnmed 2003;113:660–671 (In German and French).

108.ALD (The Academy of Laser Dentistry). Featured wavelength: diode – the diode laser in dentistry (Academy report) Wavelengths 2000:8:13.

109.Kreisler M, Meyer C, Stender E, Daubla¨nder M, Willershausen Zo¨nnchen B, d’Hoedt B. Effect of diode laser irradiation on the attachment rate of periodontal ligament cells: an in vitro study. J Periodontol 2001:72: 1312–1317.

110.Kreisler M, Al Haj H, Daublander M, Gotz H, Duschner H, Willershausen B, D’Hoedt B. Effect of diode laser irradiation on root surfaces in vitro. J Clin Laser Med Surg 2002:20:63–69.

111.Schwarz F, Sculean A, Berakdar M, Georg T, Becker J. In vivo and in vitro effects of an Er:YAG laser, a GaAlAs diode laser and scaling and root planing on periodontally diseased root surfaces. A comparative histologic study. Lasers Surg Med 2003:32:359–366.

112.Moritz A, Gutknecht N, Doertbudak O, Goharkhay K, Schoop U, Schauer P, Sperr W. Bacterial reduction in periodontal pockets through irradiation with a diode laser: a pilot study. J Clin Laser Med Surg 1997:15: 33–37.

113.Moritz A, Schoop U, Goharkhay K, Schauer P, Doertbudak O, Wernisch J, Sperr W. Treatment of periodontal pockets with a diode laser. Lasers Surg Med 1998:22:302–311.

114.Coluzzi DJ. Lasers and soft tissue curettage: An update. Compendium 2002:23:1104–1111.

115.Hibst R, Paulus R. A New approach to fluorescence spectroscopy for caries detection. Proc SPIE 1999:3593:141–147.

116.Hibst R, Paulus R, Lussi A. Detection of occlusal caries by laser fluorescence: basic and clinical investigations. Med Laser Appl 2001:16:205–213.

117.Krause F, Braun A, Frentzen M. The possibility of detecting subgingival calculus by laser-fluorescence in vitro. Lasers Med Sci 2003:18:32–35.

118.Sulewski JG. Historical survey of laser dentistry. Dent Clin North Am 2000:44:717–752.

119.Henry CA, Judy M, Dyer B, Wagner M, Matthews JL. Sensitivity of Porphyromonas and Prevotella species in liquid media to argon laser. Photochem Photobiol 1995:61:410–413.

120.Henry CA, Dyer B, Wagner M, Judy M, Matthews JL. Phototoxicity of argon laser irradiation on biofilms of Porphyromonas and Prevotella species. J Photochem Photobiol B 1996:34:123–128.

121.Finkbeiner RL. The results of 1328 periodontal pockets treated with the argon laser: selective pocket thermolysis. J Clin Laser Med Surg 1995:13:273–281.

122.Rechmann P, Henning T, Reichart P. Periodontal treatment with the frequency doubled alexandrite laser in dogs. Proc SPIE 2000:3910:351–341.

123.Frentzen M, Koort HJ, Thiensiri I. Excimer lasers in dentistry: future possibilities with advanced technology. Quintessence Int 1992:23:117–133.

124.Folwaczny M, Mehl A, Haffner C, Hickel R. Substance removal on teeth with and without calculus using 308 nm XeCl excimer laser radiation. An in vitro investigation. J Clin Periodontol 1999:26:306–312.

125.Dental Clinics of North America 48 (2004);Lasers in dental implantology : 999-1015.

126. Point of care:Journal of the Canadian Dental Association February 2005;71, 2.

127.Gerschman JA, Ruben J, Gebart-Eaglemont J. Low level laser therapy for dentinal tooth hypersensitivity. Australian Dent J 1994;39:353-7.

128.L. J. Walsh : The current status of low level laser therapy in dentistry. Australian Dental Journal 1997;42:(5):302-6.

129.Proc Natl Acad Sci U S A. 2003 Mar18; 100(6): 3439-44. Epub 2003.

130.Fujimaki Y et al. Low-level laser irradiation attenuates production of reactive oxygen species by human neutrophils. J Clin Laser Med Surg. 2003 Jun;21(3):165-70.

131.J. Ty Hopkins et al. Low-Level Laser Therapy Facilitates Superficial Wound Healing in Humans: A Triple-Blind, Sham-Controlled Study Journal of Athletic Training 2004;39(3):223–229 by the National Athletic Trainers’ Association, Inc www.journalofathletictraining.org

132.Jose Claudio et al. Clinical Study of the Gingiva Healing after Gingivectomy and Low-Level Laser Therapy; Photomedicine and Laser Surgery. October 1, 2006, 24(5): 588-594. doi:10.1089/pho.2006.24.588.

133.Juliano Milanezi et al. In Vivo Effect of Photodynamic Therapy on Periodontal Bone Loss in Dental Furcations. Journal of Periodontology 2008, Vol. 79, No. 6, Pages 1081-1088 (doi:10.1902/jop.2008.070456).

134.At http://www.wtia.com.au. WTIA-Panel 4 – Guidance Note 9 Laser safety – April 2003.